# Objective

This process flow describes how to fabricate a bilayer resist stack for lift-off of patterns defined by e-beam lithography and metallisation. The resist stack combines a lift-off resist LOR5A from Microchem (http://microchem.com/pdf/PMGI-Resists-data-sheetV-rhcedit-102206.pdf) and a standard e-beam resist CSAR from AllResist (http://www.allresist.com/csar-62-ar-p-6200/). LOR 5A spin coats to a thickness of around 300 nm at 5000 rpm and is thus suitable for undercuts of around 300 nm. If a smaller pattern dimension or higher density of patterns is needed, a thinner LOR is required. For undercuts in range of 20 – 50 nm, SF3 or LOR 1A should be used.

| ## Step Heading | ## Equipment |  | ## Comments |
| --- | --- | --- | --- |
| - Pretreatment |  |  |  |
| Surface treatment | Hotplate | Dehydrate wafers 5 min @ 210 oC |  |
| Spin coat of LOR and CSAR |  |  |  |
| Coat wafers | Spin Coater: Manual all resists (A-5) | Resist: LOR 5A Spin coating: 5 sec @ 1000 rpm, acc 200 rpm/min/s 60 sec @ 5000 rpm (for appr. 300 nm) Softbake: 10 min @ 210 oC (Hotplate set temperature: 230 oC) Resist: CSAR:Anisole 1:1 (AR-P 6200, AllResist) Spin coating: 60 sec @ 4000 rpm (for appr. 60 nm) Softbake: 1 min @ 150 oC (Hotplate set temperature: 170 oC) | Use E-beam/Anisole process bowl in A-5. It is not allowed to use ‘Spin Coater: Manual Standard Resists’ in E-5. Pre-spin of LOR5A is necessary to obtain uniform coating. Undercut etch rate depends strongly on softbake time and temperature. Use E-beam/Anisole process bowl. Clean bowl and spin coater with 1165 Remover. |
| E-beam exposure |  |  |  |
| E-beam exposure | E-beam writer | Dose: 200 - 350 µC/cm2; dose depends on resist thickness, pattern density, and developer. A dose-test is required. | Lines 200 – 500 nm |
| Development of CSAR |  |  |  |
| Develop-ment | Beaker (E-beam fumehood) | CSAR: Develop with AR 600-546, 60 sec Rinse in IPA, 60 sec, blow dry with N2 |  |
| Development of LOR5A |  |  |  |
| Develop-ment | Developer: TMAH Manual | LOR5A: Recipe: SP 15s | The developer installed on ‘Developer: TMAH Manual’ is AZ 726 MIF, which is a toxic chemical. |
| Recipe: SP 30s; Recipe: SP 15s. 400 nm undercut in LOR5A; 200 nm undercut in LOR5A 1 µm. In order to decrease the etch rate of LOR5A in AZ 726 MIF, the softbake temperature and time of LOR5A should be larger, see http://microchem.com/pdf/PMGI-Resists-data-sheetV-rhcedit-102206.pdf. |
| Metallization |  |  |  |
| Evaporation | Alcatel | E-beam evaporation of 5 nm Cr 50 nm Au Rates 10Å/s |  |
| Lift-off |  |  |  |
| Lift-off | Fumehood, D-3 | Lift-off: Beaker with Remover 1165. If possible, face resist surface downwards to avoid re-deposition of metal onto substrate surface. At 60 °C, the structures lift-off almost immediately, whereas regions without patterns took up to 30 minutes to lift. Rinse: IPA General info on lift-off of LOR: LOR and acetone creates solids. Liftoff and stripping of LOR must be done with NMP or DMSO based strippers only (http://www.nnf.ncsu.edu/processes/lithographic/microchem-lift-resist-lor-process). At high soft bake temperatures, the removal rate of LOR 5A decreases, see http://microchem.com/pdf/PMGI-Resists-data-sheetV-rhcedit-102206.pdf . | Wear 4H gloves and apron when handling Remover 1165. Lift-off is enhanced by heating up Remover 1165: This may be done on a hotplate in a fumehood. The temperature of Remover 1165 may not exceed 65 °C. Use only hotplate with an external temperature sensor to heat Remover 1165, and never leave the setup. Remover 1165 is disposed of in c-waste when cooled to room temperature. |
| SEM Inspection |  |  |  |
| - SEM Inspection | SEM Supra 2 | Inspection of metal lift off. WD 4 mm HV = 10 kV InLens detector | Due to undercut of resist, the line edge roughness of the deposited metal is among other things defined by deposition rate and substrate temperature during deposition. |