## Process and measurement data
| Aceptence test notes |  |  | Cryofox | [7 empty] | 6" wafer
10 mm
| 2010-11-11 | [10 empty] | 20 mm
30 mm
[10 empty] | 40,-40 |  | 40 mm |  | 40, 40
50 mm
60 mm
[9 empty] | 0,-70 |  |  | 70 mm |  |  | 0.7
80 mm
90 mm
[10 empty] | 100,-40 |  | 100 mm |  | 100; 40 mm
110 mm
120 mm
130 mm
140 mm
1
| Process |  |  |  |  | Start pump | 2011-05-18 | 9:33
| E-beam 1 | Ag | Film 1 |  |  | Deponering færdig |  | 10:45
| Batch | 1 | Deposition | 1
| Mask | Single E-beam
Parameters
|  |  | Thickness | 80nm |  |  | Pressure | Base | 9e-07 |  | General | Cryo proc | 10%
|  |  | Tooling | 10.9 | [7 empty] | Gas preset | 50%
| Position 1 | [5 empty] | Position 2
| Wafer 3 | Box Nora |  |  |  |  | Wafer 4 | Box Nora
| 1 | 69.71 |  | nm |  |  | 1 | 71.63
| 2 | 71.1 | 70.63 |  |  |  | 2 | 68.96
| 3 | 70.09 | 70.4 |  |  |  | 3 | 69.34
| 4 | 69.16 | 71.73 |  |  |  | 4 | 68.09
| 5 | 69.45 | 70.35 |  |  |  | 5 | 67.44
| 6 | 68.96 | 69.45 |  |  |  | 6 | 67.89
| 7 | 68.18 | 68.51 |  |  |  | 7 | 67.53
| 8 |  | 68.25 |  |  |  | 8 | 67.77
| 9 | [5 empty] | 9 | 67.62
| Average | 69.5214 | 69.9029 |  |  |  | Average | 68.4744
| Min |  | 68.25 |  |  |  | Min | 67.44
| Max |  | 71.73 |  |  |  | Max | 71.63
|  |  | -2.3645 |  |  |  |  | -1.5107
|  |  | 2.6138 |  |  |  |  | 4.6084
2
| Process | 2
| E-beam 1 | Ag | Film 1
| Batch | 2 | Deposition | 2
| Mask | Single E-beam
Parameters
|  |  | Thickness | 80nm |  |  | Pressure | Base | 9e-07 |  | General | Cryo proc | 10%
|  |  | Tooling | 10.9 | [7 empty] | Gas preset | 50%
| Position 1 | [5 empty] | Position 2
| Wafer 5 | Box Nora |  |  |  |  | Wafer 6 | Box Nora
| 1 | 67.43 |  |  |  |  | 1 | 66.99
| 2 | 66.88 |  |  |  |  | 2 | 67.7
| 3 | 68.24 |  |  |  |  | 3 | 67.44
| 4 | 68.05 |  |  |  |  | 4 | 68.14
| 5 | 68.98 |  |  |  |  | 5 | 67.65
| 6 | 68.32 |  |  |  |  | 6 | 66.24
| 7 | 66.92 |  |  |  |  | 7 | 66.95
| 8 | 66.97 |  |  |  |  | 8 | 67.14
| 9 | 69.04 |  |  |  |  | 9 | 67.48
| Average | 67.87 |  |  |  |  | Average | 67.3033
| Min | 66.88 |  |  |  |  | Min | 66.24
| Max | 69.04 |  |  |  |  | Max | 68.14
|  | -1.4587 | [5 empty] | -1.5799
|  | 1.7239 | [5 empty] | 1.2431
3 |  |  |  | 2011-05-18 | 17:00 | Running time
| Process | 3
| Co ebeam |  |  |  |  | E-beam 1 | Cr | Film 2 |  | Ebeam 2 | NI | Film 3
| Batch | 2 | Deposition | 2 |  | Tooling | 10.0 |  |  | Tooling | 10.0
| Mask | Single E-beam |  |  |  | Rate | 0.5 |  |  | Rate 0,5
Parameters
|  |  | Thickness | 80nm |  |  | Pressure | Base | 9e-07 |  | General | Cryo proc | 10%
[12 empty] | Gas preset | 50%
| Position 1 | [5 empty] | Position 2
| Wafer 7 | Box Nora |  |  |  |  | Wafer 8 | Box Nora
| 1 | 80.41 | NM
| 2 | 79.84 |  | Cu in the film
| 3 | 79.12
| 4 | 80.39
| 5 | 78.87
| 6 | 78.62
| 7 | 78.5 |  |  |  |  | XPS: | Cu found in the film (wafer 7)
| 8 | 78.12
| 9 | 76.54
| Average | 78.9344
| Min | 76.54
| Max | 80.41
-3.0335
1.8693
4
| Process | 4
| E-beam 1 | Ag | Film 1
| Batch | 3 | Deposition | 3
| Mask | Single E-beam
Parameters
|  |  | Thickness | 80nm |  |  | Pressure | Base | 9e-07 |  | General | Cryo proc | 10%
|  |  | Tooling | 10.9 | [7 empty] | Gas preset | 50%
| Position 1 | [5 empty] | Position 2
| Wafer9 | Box Nora |  |  |  |  | Wafer 10 | Box Nora
| 1 | 65.3 |  |  |  |  | 1 | 65.02
| 2 | 68.82 |  |  |  |  | 2 | 66.52
| 3 | 65.03 |  |  |  |  | 3 | 65.69
| 4 | 65.17 |  |  |  |  | 4 | 65.28
| 5 | 66.82 |  |  |  |  | 5 | 67.41
| 6 | 65.79 |  |  |  |  | 6 | 65.04
| 7 | 67.76 |  |  |  |  | 7 | 66.67
| 8 | [5 empty] | 8 | 67.1
| 9 | [5 empty] | 9 | 65.25
| Average | 66.3843 |  |  |  |  | Average | 65.9978
| Min | 65.03 |  |  |  |  | Min | 65.02
| Max | 68.82 |  |  |  |  | Max | 67.41
|  | -2.0401 | [5 empty] | -1.4815
|  | 3.6691 | [5 empty] | 2.1398
5 |  |  |  | 2011-05-19 | 8:00 | Running time
| Process | [15 empty] | Tooling 18/5
| Co ebeam |  |  |  |  | E-beam 1 | Cr | Film 2 |  | Ebeam 2 | NI | Film 3 |  |  |  | E-beam 1 | Cr | Film 2 | Ebeam 2 | NI | Film 3
| Batch | 2 | Deposition | 2 |  | Tooling | 10.0 |  |  | Tooling | 10.0 |  |  |  |  | Tooling | 11.4 |  | Tooling | 10.6
| Mask | Single E-beam |  |  |  | Rate | 0.5 |  |  | Rate 0,5 | [5 empty] | Rate | 0.5 |  | Rate 0,5
Parameters
|  |  | Thickness | 80nm |  |  | Pressure | Base | 9e-07 |  | General | Cryo proc | 10%
[12 empty] | Gas preset | 50%
| Position 1 | [5 empty] | Position 2
| Wafer 11 | Box Nora |  |  |  |  | Wafer 12 | Box Nora
| 1 | 83.54 |  |  |  |  | 1 | 80.79
| 2 | 82.92 |  |  |  |  | 2 | 82.15
| 3 | 79.97 |  |  |  |  | 3 | 81.11
| 4 | 80.05 |  |  |  |  | 4 | 78.03
| 5 | 79.47 |  |  |  |  | 5 | 78.81
| 6 | 79.16 | 78.39 |  |  |  | 6 | 77.82
| 7 | 79.64 |  |  |  |  | 7 | 78.8
| 8 | 80.48 |  |  |  |  | 8 | 77.64
| 9 | 79.83 |  |  |  |  | 9 | 78.42
| average | 80.5622 |  |  |  |  | Average | 79.2856
| min | 79.16 |  |  |  |  | Min | 77.64
| max | 83.54 |  |  |  |  | Max | 82.15
| min | -1.7405 | [5 empty] | -2.0755
| max | 3.6962 | [5 empty] | 3.6128
| Process | [5 empty] | 2011-05-19 | 9:24
| E-beam 1 | Ag | Film 1
| Batch | 4 | Deposition
| Mask | Single E-beam
Parameters
|  |  | Thickness | 80nm |  |  | Pressure | Base | 9e-07 |  | General | Cryo proc | 10%
|  |  | Tooling | 10.9 | [7 empty] | Gas preset | 50%
| Position 1 | [5 empty] | Position 2
| Wafer 13 | Box Nora |  |  |  |  | Wafer 14 | Box Nora
| 1 | 71.42 |  |  |  |  | 1 | 65.66 | 65.57
| 2 | 68.22 |  |  |  |  | 2 | 67.74 | 68.21
| 3 | 67.58 |  |  |  |  | 3 |  | 66.09
| 4 | 65.71 |  |  |  |  | 4 |  | 67.03
| 5 | 67.64 |  |  |  |  | 5 |  | 66.18
| 6 | 66.77 |  |  |  |  | 6 |  | 66.46
| 7 | 68.13 |  |  |  |  | 7 |  | 68
| 8 | 66.78 |  |  |  |  | 8 |  | 67.86
| 9 | 66.17 |  |  |  |  | 9 |  | 67.05
| average | 67.6022 |  |  |  |  | average | 66.7 | 66.9389
| min | 65.71 |  |  |  |  | min | 65.66 | 65.57
| max | 71.42 |  |  |  |  | max | 67.74 | 68.21
| min | -2.7991 |  |  |  |  | min | -1.5592 | -2.045
| max | 5.6474 |  |  |  |  | max | 1.5592 | 1.8989
5 |  |  |  | 2011-05-19 | 10:53 | Pump down large chamber
Process
| Co ebeam |  |  |  |  | E-beam 1 | Cr | Film 2 |  | Ebeam 2 | NI | Film 3
| Batch | 2 | Deposition | 2 |  | Tooling | 10.0 |  |  | Tooling | 10.0
| Mask | Single E-beam |  |  |  | Rate | 0.5 |  |  | Rate 0,5
Parameters
|  |  | Thickness | 80nm |  |  | Pressure | Base | 9e-07 |  | General | Cryo proc | 10%
[12 empty] | Gas preset | 50%
| Position 1 | [5 empty] | Position 2
| Wafer 17 | Box Nora |  |  |  |  | Wafer 18 | Box Nora | Wafer placed in position 17, Measured and put into place 18!!
| 1 | 82.7 |  |  |  |  | 1 | 79.31
| 2 | 79.91 |  |  |  |  | 2 | 80.66
| 3 | 80.37 |  |  |  |  | 3 | 79.98
| 4 | 79.66 |  |  |  |  | 4 | 78.95
| 5 | 78.65 |  |  |  |  | 5 | 80.02
| 6 | 79.43 | 76.16 |  |  |  | 6 | 78.17
| 7 | 79.61 |  |  |  |  | 7 | 79.51
| 8 | 78.33 |  |  |  |  | 8 | 79.7
| 9 | 79.99 |  |  |  |  | 9 | 78.89
| Average | 79.85 |  |  |  |  | Average | 79.4656
| Min | 78.33 |  |  |  |  | Min | 78.17
| Max | 82.7 |  |  |  |  | Max | 80.66
| min | -1.9036 |  |  |  |  | min | -1.6303
| max | 3.5692 |  |  |  |  | max | 1.5031
[5 empty] | 2011-05-19 | 14:23
Process
|  | Co ebeam |  |  |  |  | E-beam 1 | Cr | Film 2 |  | Ebeam 2 | NI | Film 3
|  | Batch | 2 | Deposition | 2 |  | Tooling | 10.0 |  |  | Tooling | 10.0
|  | Mask | Single E-beam |  |  |  | Rate | 0.5 |  |  | Rate 0,5
Parameters
|  |  |  | Thickness | 80nm |  |  | Pressure | Base | 9e-07 |  | General | Cryo proc | 10%
[13 empty] | Gas preset | 50%
Stopped in the middle of process, cleaned chamber afterwards
|  | Position 1 | [5 empty] | Position 2
|  | Wafer 19 | Box Nora |  |  |  |  | Wafer 20 | Box Nora
[5 empty] | 2011-05-19 | 14:23
Process
|  | Co ebeam |  |  |  |  | E-beam 1 | Cr | Film 2 |  | Ebeam 2 | NI | Film 3
|  | Batch | 2 | Deposition | 2 |  | Tooling | 10.0 |  |  | Tooling | 10.0
|  | Mask | Single E-beam |  |  |  | Rate | 0.5 |  |  | Rate 0,5
Parameters
|  |  |  | Thickness | 80nm |  |  | Pressure | Base | 9e-07 |  | General | Cryo proc | 10%
[13 empty] | Gas preset | 50%
|  | Position 1 | [5 empty] | Position 2 |  |  |  |  | Wafer 22 |  | Large center structure: |  | 80.72 nm
|  | Wafer 21 | Box Nora |  |  |  |  | Wafer 22 | Box Nora |  |  |  | Again |  |  |  | 81.97
|  | 1 | 80.44 |  |  |  |  | 1 | 85.13 |  |  |  | 1
|  | 2 | 78.07 |  |  |  |  | 2 | 81.53 |  |  |  | 2
|  | 3 | 80.01 |  |  |  |  | 3 | 81.77 |  |  |  | 3
|  | 4 | 79.52 |  |  |  |  | 4 | 80.62 |  |  |  | 4 | 80.48
|  | 5 | 80.63 |  |  |  |  | 5 | 81.29 |  |  |  | 5 | 80.91
|  | 6 | 80.74 |  |  |  |  | 6 | 81.43 |  |  |  | 6 | 81.55
|  | 7 | 80.95 |  |  |  |  | 7 | 82.62 |  | 81.74 |  | 7
|  | 8 | 78.8 |  |  |  |  | 8 | 82.34 |  |  |  | 8
|  | 9 | 80.79 |  |  |  |  | 9 | 81.44 |  |  |  | 9
|  | Average | 79.9944 | nm |  |  |  | Average | 82.0189 | nm |  |  | Average | 80.98
|  | Min | 78.07 |  |  |  |  | Min | 80.62 |  |  |  | Min | 80.48
|  | Max | 80.95 |  |  |  |  | Max | 85.13 |  |  |  | Max | 81.55
|  |  | -2.4057 | [5 empty] | -1.7056 |  |  |  |  | -0.61744
|  |  | 1.1945 | [5 empty] | 3.7932 |  |  |  |  | 0.70388

## Summary NiCr
Summary of depositions test: NiCr deposition
The table below shows the NiCr thickness measured at 9 positions on each wafer. The 9 points are evenly distributed aling a line on the wafers. During the deposition the line was placed along the radius of the wafer holder tool in the machine.
Average thickness is calcultated below.
Deposition parameters
NiCr |  | Batch 1 |  | Batch 2 |  | Batch 3
|  | Wafer 11 | Wafer 12 | Wafer 17 | Wafer 18 | Wafer 21 | Wafer 22 | Wafer Box "Nora" | Wafer 11*, 12 (batch 1)
Thickness | Point 1 | 83.54 | 80.79 | 82.7 | 79.31 | 80.44 | 85.13 | nm | Mask | Single E-beam
| 2 | 82.92 | 82.15 | 79.91 | 80.66 | 78.07 | 81.53 |  | Ebeam 1 |  | Ebeam 2
| 3 | 79.97 | 81.11 | 80.37 | 79.98 | 80.01 | 81.77 |  | Material | Cr | Material | Ni
| 4 | 80.05 | 78.03 | 79.66 | 78.95 | 79.52 | 80.62 |  | Film number | 2 | Film number | 3
| 5 | 79.47 | 78.81 | 78.65 | 80.02 | 80.63 | 81.29 |  | Tooling | 11.4 | Tooling | 10.6
| 6 | 79.16 | 77.82 | 79.43 | 78.17 | 80.74 | 81.43 |  | Rate (Å/s) | 0.5 | Rate (Å/s) | 0.5
| 7 | 79.64 | 78.8 | 79.61 | 79.51 | 80.95 | 82.62
| 8 | 80.48 | 77.64 | 78.33 | 79.7 | 78.8 | 82.34
| 9 | 79.83 | 78.42 | 79.99 | 78.89 | 80.79 | 81.44 |  | Wafer 17*, 18 (batch 2)
[9 empty] | Mask | Single E-beam
Single wafer | Average thickness | 80.56 | 79.29 | 79.85 | 79.47 | 79.99 | 82.02 | nm | Ebeam 1 |  | Ebeam 2
| Min | -1.74 | -2.08 | -1.9 | -1.64 | -2.4 | -1.71 |  | Material | Cr | Material | Ni
| Max | 3.7 | 3.61 | 3.57 | 1.5 | 1.2 | 3.79 |  | Film number | 2 | Film number | 3
[9 empty] | Tooling | 10 | Tooling | 10
Wafer to wafer | Average | 79.92 |  | 79.6578 |  | 81.0067 |  |  | Rate (Å/s) | 0.5 | Rate (Å/s) | 0.5
| Min | -2.85 |  | -1.87 |  | -3.63
| Max | 4.53 |  | 3.82 |  | 5.09
Wafer 21, 22 (batch 3)
[9 empty] | Mask | Single E-beam
Batch to batch | Average | 80.2 | [6 empty] | Ebeam 1 |  | Ebeam 2
| Min | -3.19 | [6 empty] | Material | Cr | Material | Ni
| Max | 6.15 | [6 empty] | Film number | 2 | Film number | 3
[9 empty] | Tooling | 10 | Tooling | 10
| Min Cr At. % | (chosen; larger etch timeinterval) |  |  | 33.79 |  | Min Ni At. % | 47.59 | Rate (Å/s) | 0.5 | Rate (Å/s) | 0.5
| Max Cr At % |  |  |  | 44.6 |  | Max Ni At. % | 55.87
* = checked with XPS. See result in LabAdviser
| Min Cr At % | (chosen; smaller etch time interval) |  |  | 34.67 |  | Min Ni At. % | 47.59
| Max Cr At % |  |  |  | 43.97 |  | Max Ni At. % | 55.87

## Summary Ag
Summary of depositions test: Ag deposition
The table below shows the Ag thickness measured at 9 positions on each wafer. The 9 points are evenly distributed aling a line on the wafers. During the deposition the line was placed along the radius of the wafer holder tool in the machine.
Average thickness is calcultated below.
Deposition parameters
Ag |  | Batch 1 |  | Batch 2 |  | Batch 3 |  | Batch 4
|  | Wafer 3 | Wafer 4 | Wafer 5 | Wafer 6 | Wafer9 | Wafer 10 | Wafer 13 | Wafer 14 | Wafer Box "Nora"
Thickness | Point 1 | 69.71 | 71.63 | 67.43 | 66.99 | 65.3 | 65.02 | 71.42 | 65.57 |  | Batches 1,2,3,4
| 2 | 71.1 | 68.96 | 66.88 | 67.7 | 68.82 | 66.52 | 68.22 | 68.21 |  | Mask | Single E-beam
| 3 | 70.09 | 69.34 | 68.24 | 67.44 | 65.03 | 65.69 | 67.58 | 66.09 |  | Source | Ebeam 1
| 4 | 69.16 | 68.09 | 68.05 | 68.14 | 65.17 | 65.28 | 65.71 | 67.03 |  | Material | Ag
| 5 | 69.45 | 67.44 | 68.98 | 67.65 | 66.82 | 67.41 | 67.64 | 66.18 |  | Film number | 2
| 6 | 68.96 | 67.89 | 68.32 | 66.24 | 65.79 | 65.04 | 66.77 | 66.46 |  | Tooling | 10.9
| 7 | 68.18 | 67.53 | 66.92 | 66.95 | 67.76 | 66.67 | 68.13 | 68
| 8 | 68.25 | 67.77 | 66.97 | 67.14 |  | 67.1 | 66.78 | 67.86
| 9 |  | 67.62 | 69.04 | 67.48 |  | 65.25 | 66.17 | 67.05
Single wafer | Average | 69.36 | 68.47 | 67.87 | 67.3 | 66.38 | 66 | 67.6 | 66.94
| Min | -1.7 | -1.5 | -1.46 | -1.58 | -2.03 | -1.48 | -2.8 | -2.05
| Max | 2.51 | 4.62 | 1.72 | 1.25 | 3.68 | 2.14 | 5.65 | 1.9
Wafer to wafer | Average | Wafer 3: 68.89; Wafer 5: 67.5867; Wafer9: 66.17; Wafer 13: 67.27
| Min | Wafer 3: -2.1; Wafer 5: -1.99; Wafer9: -1.74; Wafer 13: -2.53
| Max | Wafer 3: 3.98; Wafer 5: 2.15; Wafer9: 4; Wafer 13: 6.17
Batch to batch | Average | Wafer 3: 67.5
| Min | Wafer 3: -3.67; Wafer9: 71.63; Wafer 10: 68.325; Wafer 14: 0.096744
| Max | Wafer 3: 6.12; Wafer9: 65.02; Wafer 10: 6.61