## Sheet1
Mini study: |  |  |  | Stress in thin Ni films |  |  | Performed January-April 2014 |  |  | Katharina Nilson
Tested: |  | Stress in Ni films deposited in Wordentec
Deposition rate was changed to see if that influenced the stress in the films
Depositions: |  | Deposition of 200 nm Ni with different deposition rates.
In two cases the film thickenss was 50 and 110 nm instead.
Measurements: |  | Wafer thickness was measured with thickness measurer
Stress was measured with Dektak 8
Details Dektak 8: |  | The film thickness was set to 200 nm in all calculations, with two exceptions (when 50 and 110 nm was deposited).
Thickness of substrate were set to 525 micrometer for the first three depositions.
The Dektak measurements were leveled in the start and end position of the scan.
Program "stress 4 inch" was used (saved under data/knil/Ni stress) with parameters: range= 1mm, length= 80000my, time=60 s.
The stress was measued across the wafer along two perpendicular lines (1st and 2nd value).
Explanations to the numbers below:
| 1. | When measuring the stress with the Dektak 8 it gives average compressiv and tensile stress values for along the line scan. These numbes are written under "Flat right (1st measurement)"
| 2. | After the first measurement the wafer is rotated and the measurement is done along a line perpedicular to the first one. These numbers are written under "Flat back (2nd measurement)"
| 3. | An average value are calcuated for each wafer (from the 1st and 2nd measurements). These are written under "Average stress ( 1st and 2nd measurements)"
| 4. | For each deposition rate, two or three film were deposited. An average for these are calcultated in the and written under "Average stress calculated from the different wafers."- the fields are marked green below.
Measurements
[6 empty] | Calculated stress (Mpa) |  |  |  | Average stress ( 1st and 2nd measurements)
[6 empty] | Flat right (1st measurement) |  | Flat back (2nd measurement)
Name of Wafer | Wafer thickness | Deposition date | Thickness | Deposition rate | Measurement date | Compressive average | Tensile average | Compressive average | Tensile average | Comp | Tensile | Power soak 2 | Actual rate | Pressure
Wafer 19 |  | 2014-01-08 | 200 nm | 15 Å/s | 2014-01-08 | 0 | 972 | 0 | 956 | 0 | 964 | 47 | 15 | 3.3e-06
Wafer 20 |  | 2014-01-08 | 200 nm | 7 Å/s | 2014-01-08 | 0 | 901 | 0 | 873 | 0 | 887 | 36 | 7 | 3.3e-06
Wafer 21 |  | 2014-01-08 | 190 nm | 1 Å/s | 2014-01-08 | 0 | 365 | 0 | 328 | 0 | 346.5 | 20 | 1 | 2.5e-06
Wafer 25 | 0.5256 | 2014-01-22 | 200 nm | 3Å/s | 2014-01-22 | 0 | 763 | 0 | 759 | 0 | 761 | 30 | 3 | P=2.1E-6
[6 empty] | flat right |  | flat back
20140319 |  | 2014-03-19 10:30 | 200 nm | 15 Å/s | 2014-03-19 10:30 | 0 | 1157 | 0 | 1090 | 0 | 1123.5 | 41% | 13-16 | 3e-06
Measurements to check time dependence |  |  |  |  | 2014-03-20 | 0 | 821 | 0 | 771 | 0 | 796
[5 empty] | 2014-03-21 | 0 | 580 | 0 | 563 | 0 | 571.5
[5 empty] | 2014-03-25 | 0 | 348 | 1.1 | 338 | 0.55 | 343
[5 empty] | 2014-04-01 | 0 | 324 | 0 | 307 | 0 | 315.5
[5 empty] | 2014-04-04 | 0 | 334 | 2.9 | 288 | 1.45 | 311
[6 empty] | flat right |  | flat back
20140401 W9 | 0.5245 | 2014-04-01 | 200 nm | 15 Å/s | 2014-04-01 | 0 | 1060 | 0 | 1045 | 0 | 1052.5 | 42 | 15-16 Å/s | p=3.2E-6
Measurements to check time dependence |  |  |  |  | 2014-04-04 | 0 | 495 | 0 | 493 | 0 | 494
[5 empty] | 2014-04-10 | 0 | 329 | 0 | 337 | 0 | 333
20140401 W10 | 0.5245 | 2014-04-01 | 200 nm | 3 Å/s | 2014-04-01 | 0 | 860 | 0 | 834 | 0 | 847 | 30 | 3-3,5 Å/s | p=3.5E-6
[6 empty] | flat right |  | flat back
20140408 W5 | 0.5163 | 2014-04-08 | 50 nm | 1 Å/s | 2014-04-08 | 22 | 520 | 0 | 367 | 11 | 443.5 | Set to 23 % (far to low) | 0.1-1 | p=3,3E-6
20140408 W6 | 0.5167 | 2014-04-08 | 110 nm | 3 Å/s | 2014-04-08 | 0 | 602 | 3 | 531 | 1.5 | 566.5 | 35% | 3Å/s | p=4E-6
20140408 W7 | 0.516 | 2014-04-08 | 200 nm | 10 Å/s | 2014-04-08 | 0 | 934 | 0 | 938 | 0 | 936 | 40.5% | 7-10Å/s | p=4E-6
20140417 W3 | 0.5235 | 2014-04-16 | 200 nm | 7 Å/s | 2014-04-16 | 0 | 925 | 0 | 913 | 0 | 919 | 36.5% | 7 Å/s | p=3,5E-6
20140416 W 4 | 0.523 | 2014-04-16 | 200 nm | 10 Å/s | 2014-04-16 | 0 | 951 | 0 | 938 | 0 | 944.5 | 40.5%. | 12-10 Å/s | p=2.8E-6
Analysis of the numbers above
Stress depending on the deposition rate
|  | Stress (Mpa) measured along two lines on the wafer, perpendicular to each other. |  |  |  | Average stress (calculated from 1st and 2nd measurement) |  | Average stress calculated from the different wafers. |  |  |  |  | Saok power 2 settings for desired rate
|  | Flat right (1st measurement) |  | Flat back (2nd measurement) | [8 empty] | With the correct soak power settings the rate will be correct from the beginning of the process.
Deposition rate (Å/s) | Deposition ane measurment date | Compressive average | Tensile average | Compressive average | Tensile aver | Comp | Tensile | Comp | Tensile | Standard deviation Tensile stress |  |  | Rate (Å/s) | Recommended soak power 2 setting (%)
[13 empty] | 1 | 25
15 | 2014-03-19 10:30 | 0 | 1157 | 0 | 1090 | 0 | 1123.5 | 0 | 1046.6667 | 79.9098 |  |  | 3 | 32
15 | 2014-04-01 | 0 | 1060 | 0 | 1045 | 0 | 1052.5 | [5 empty] | 7 | 37
15 | 2014-01-08 | 0 | 972 | 0 | 956 | 0 | 964 | [5 empty] | 10 | 40.5
[13 empty] | 15 | 43
10 | 2014-04-08 | 0 | 934 | 0 | 938 | 0 | 936 | 0 | 940.25 | 6.0104
10 | 2014-04-16 | 0 | 951 | 0 | 938 | 0 | 944.5
7 | 2014-01-08 | 0 | 901 | 0 | 873 | 0 | 887 | 0 | 903 | 22.6274
7 | 2014-04-16 | 0 | 925 | 0 | 913 | 0 | 919
3 | 2014-04-08 | 0 | 602 | 3 | 531 | 1.5 | 566.5 | 0.5 | 724.8333 | 143.7048 |  |  | Time dependence:
3 | 2014-01-22 | 0 | 763 | 0 | 759 | 0 | 761
3 | 2014-04-01 | 0 | 860 | 0 | 834 | 0 | 847 | [5 empty] | Wafer | Deposition date | Deposition rate | Measurement date | Compressive average | Tensile average | Compressive average | Tensile aver | Comp | Tensile | Number of days between deposition and measurement
[13 empty] | Wafer 1 deposited with 15 Å/s | 2014-03-19 10:30 | 15 Å/s | 2014-03-19 10:30 | 0 | 1157 | 0 | 1090 | 0 | 1123.5 | 0
1 | 2014-04-08 | 22 | 520 | 0 | 367 | 11 | 443.5 | 5.5 | 395 | 68.5894 | [5 empty] | 2014-03-20 | 0 | 821 | 0 | 771 | 0 | 796 | 1
1 | 2014-01-08 | 0 | 365 | 0 | 328 | 0 | 346.5 | [8 empty] | 2014-03-21 | 0 | 580 | 0 | 563 | 0 | 571.5 | 2
[16 empty] | 2014-03-25 | 0 | 348 | 1.1 | 338 | 0.55 | 343 | 6
[16 empty] | 2014-04-01 | 0 | 324 | 0 | 307 | 0 | 315.5 | 13
[16 empty] | 2014-04-04 | 0 | 334 | 2.9 | 288 | 1.45 | 311 | 16
[13 empty] | Wafer 2 deposited with 15 Å/s | 2014-04-01 | 15 Å/s | 2014-04-01 | 0 | 1060 | 0 | 1045 | 0 | 1052.5 | 0
[16 empty] | 2014-04-04 | 0 | 495 | 0 | 493 | 0 | 494 | 3
[16 empty] | 2014-04-10 | 0 | 329 | 0 | 337 | 0 | 333 | 9