## Slide 1: Dektak XT and P17 ”flat” scans
Measurements of surfaces that should be ”flat”

## Slide 2
DektakXT 6” blank Si center spot 1
Same scan, variation in leveling cursors
Range 15-20 nm

## Slide 3
DektakXT 6” blank Si center spot 1
Same scan, variation in leveling cursors
Range in one case 20 nm
In the other 6 nm

## Slide 4: DektakXT 6” blank Si near edge spot 1 2 separate scans Jun 23, 2025 app. 15 nm range variation scan-scan

## Slide 5
DektakXT 6” blank Si near edge spot 2
Range app 20 nm
Variation scan-scan

## Slide 6
DektakXT 6” blank Si near edge spot 3
Range app 10 nm
Variation scan-scan

## Slide 7
DektakXT 6” blank Si near edge spot 4
Range app 40 nm

## Slide 8
DektakXT 6” blank Si near edge spot 4
Same scan as previous slide
Alternate leveling
Range 25-60 nm

## Slide 9
P17
6” blank Si wafer
near edge
Spot 1
Jun 20, 2025
Range app 10 nm

## Slide 10
P17
6” blank Si near edge
Ave of 2
Spot 2
Range app 10 nm

## Slide 11
P17
6” blank Si
near edge
Ave of 2
Spot 3
Range app 17 nm

## Slide 12
P17
6” blank Si wafer Near edge spot 4
Ave of 2 scans
Jun 20, 2025
Range app 40 nm

## Slide 13
P17
6” blank Si wafer
Center
Ave of 2-3 scans in each graph
Scans at diff speed and resolution, same location
Range app 15 nm

## Slide 14
P17
6” blank Si wafer with Ni
near edge
Ave of 2 scans
Jun 20, 2025
Range app 10 nm

## Slide 15
P17
6” blank Si wafer with 10 nm Ti/ 90 nm Ni
near center
Ave of 2 scans
Jun 20, 2025
Range app 10 nm

## Slide 16
P17
6” blank Si wafer with TiOx
Various positions
Ave of 2 scans in each graph
Jun 20, 2025