## Slide 1
ICP etching of alumina on a silicon wafer
Microscale
Nanoscale

## Slide 2
Parameter | Recipe: NA_membrane | 
Description: Etch rate is 6.25 nm/min Selectivity (mask/alumina): 12:5 |  | 
Coil Power [W] | 300 | 300
Platen Power [W] | 15 | 15
Platen temperature [oC] | 0 | 0
BCl3 flow [sccm] | 20 | 20
Cl2 flow [sccm] | 7 | 7
Pressure [mTorr] | 1.2 | 1.2
Important information | 
Lithography writing | DUV lithography- 65 nm BARC layer, 360 nm M230Y resist,
Design | Hexagonal array of circular holes (7x7mm) Holes diameter- 200 nm Pitch- 400 nm
Parameter | Recipe: MA_membrane
Description: High selectivity, 25nm/min | 
Coil Power [W] | 500
Platen Power [W] | 70
Platen temperature [oC] | 0
BCl3 flow [sccm] | 40
Cl2 flow [sccm] | 15
Pressure [mTorr] | 3
Important information | 
Lithography writing | MLA-3, 1.3 µm AZ MIR resist 701
Design | 1. Hexagonal array of circular holes (7x7mm) Holes diameter- 1 µm Pitch- 2 µm 2. Micro channel Width- 500 µm
Nanoscale
Microscale

## Slide 3
Nanoscale
(small samples on Silicon carrier wafer)
P=pressure(mTorr), Po= coil power/platen power (watts), G=gas flow (sccm)- BCl3/Cl2, T= time (s)

## Slide 4
Important information | 
Same parameters | 
Po= 300/15, G= 20/7 | 
Different parameters | 
a | P=2, T=6min
b | P=2.8, T=8.3min
b
a
P=pressure(mTorr), Po= coil power/platen power (watts), G=gas flow (sccm)- BCl3/Cl2, T= time (s)
c
d
Important information | 
Parameters | 
c,d | Po= 300/15, G= 20/7, P=2, T=6min
The process is repeated on two samples
Here 50nm thick alumina is etched in all samples. The residues on the etched surface were cleaned by immersing them in the NMP solvent for 15 minutes under sonication.

## Slide 5
Nanoscale
(on 6-inch Si wafer)
P=pressure(mTorr), Po= coil power/platen power (watts), G=gas flow (sccm)- BCl3/Cl2, T= time (s)

## Slide 6
The BARC layer is etched for 12 minutes in Pegasus 2 at low pressure
Parameters | O2 (sccm) | Platen Power (W)
BARC etch @ 20 °C | 50 | 10
BARC layer- 65nm
Resist thickness after BARC etch 229.3 nm

## Slide 7
After BARC etch of 1min at 5mTorr in ICP metal etch
285.7 nm
To test the etch profile and selectivity the BARC layer was also etched in ICP metal etch.
Parameters | O2 (sccm) | Platen Power (W)
BARC etch @ 20 °C | 10 | 120

## Slide 8
Important information
Same parameters
Po= 300/15, G= 20/7, P=1.5, T=8min
Different parameters
The etch load in ‘b’ is 50% less than ‘a’ but they have the same critical dimensions
b
a
Important information | 
Same parameters | 
Po= 300/15, G= 20/7, T=8 min | 
Different parameters | 
c | P=1.3
d | P=1.5
d
c
After BARC etch, the alumina layer is etched on a 6-inch full wafer followed by photoresist stripping in the same chamber
Here b and d are the same images

## Slide 9
Important information | 
Same parameters | 
Po= 300/15, G= 20/7 | 
Different parameters | 
a | P=2.8, T=8.3min
b | P=1.8, T=7.5min
b
a
At pressure 2.8mTorr the top part of alumina is eroded so pressure is reduced to 1.8. Even etching for more longer time than required also effects the sidewall angle

## Slide 10
Microscale
Small samples
P=pressure(mTorr), Po= coil power/platen power (watts), G=gas flow (sccm)- BCl3/Cl2, T= time (s)

## Slide 11
Important information | 
Same parameters | 
Po= 500/70, G= 40/15 | 
Different parameters | 
a | T= 250, P= 3
b | T= 200, P= 3.5
P=pressure(mTorr), Po= coil power/platen power (watts), G=gas flow (sccm)- BCl3/Cl2, T= time (s)
Important information | 
Same parameters | 
G= 40/15 | 
Different parameters | 
c | Po= 500/70, T=250s, P=3
d | Po= 450/100, T=200s, P=4
The platen power and pressure can alter the mask (alumina layer) profile. In fig d, 100W platen power eroded the top part of alumina.
b
a
d
c