## Slide 1: Thin film Tool package training (TPT) course at DTU Nanolab 1st part October 2019

## Slide 2
  Learning goals:
  Be aware of relevant material properties for thin films made in the cleanroom
  Understand importance of correct process pressure (for vacuum processes)
  Understand different thin film techniques
  Compare and be able to select appropriate thin film technique
Thin film TPT course - Learning goals and course schedule
Notes: Learning goals:
A list of what we expect you to learn from this course

## Slide 3
  Plan for 1st part of the course (Today’s plan):
  Lecture (2 hours):
  An introduction to thin films
  Material properties
  Overview of different thin film techniques at DTU Nanolab - 1st part:
    Thermal oxidation
    Electroplating
Thin film TPT course - Learning goals and course schedule
Notes: Learning goals:
A list of what we expect you to learn from this course

## Slide 4
  Plan for 2nd part of the course:
  Lecture (3 hours):
  Vacuum basics
  Overview of different thin film techniques at DTU Nanolab - 2nd part:
    PVD (physical vapour deposition):
      E-beam evaporation
      Thermal evaporation
      Sputtering
    CVD (chemical vapour deposition):
      LPCVD
      PECVD
    ALD (atomic layer deposition)
Before you use an equipment:
  Contact training@nanolab.dtu.dk for individual training sessions for each thin film equipment (usually 1-2 sessions for each piece of equipment)
  Receive user competences in LabManager
Thin film TPT course - Learning goals and course schedule
Notes: Learning goals:
A list of what we expect you to learn from this course

## Slide 5
  Introduction to thin films
  Before you start – Material properties to consider
  Overview of techniques
  Equipment overview in LabAdviser
  Thin film techniques – 1st part
    Thermal oxidation
    Electroplating
Thin film TPT course – 1st part
Notes: SEM course contents

## Slide 6
  Introduction to thin films
  Before you start – Material properties to consider
  Overview of techniques
  Equipment overview in LabAdviser
  Thin film techniques – 1st part
    Thermal oxidation
    Electroplating
Thin film TPT course – 1st part
Notes: SEM course contents

## Slide 7
Atomic separation on a substrate – not yet a thin film
Thin film: Nanometer (monolayer) to several micrometers
Special for thin films:
High surface/volume ratio
Physical properties may be different from bulk
  Electrical properties, stress, density…
Å to nm
nm to μm
“thin film”
Substrate
Silicon wafer
Other semiconductors (III-V, II-VI, etc.)
Others: Metals, glass, polymer, ceramics, …
Bulk material
Hundreds of atomic layers
Once a film is above critical thickness, it behaves essentially as a bulk material
Introduction – What is a ”thin film”?
Notes: What is a thin film and how thin can the film be so that we consider as a thin film ?
The top image showing atoming separation on a substrate and the thickness can be from A to a few nanometers so this’s too thin too be a thin film.
And when there is more material on the substrate then it will start to form a film. So a thin film can be from a monolayer to several microns or until approaching bulk material properties. Thin films have a very high surface to volume ratio.
What is so special about thin films? They can have some properties that different from bulk.
Once a film is above a critical thickness, it behaves essentially as a bulk material
Depending on the property, however, this critical thickness ranges from 100 Angstroms (e.g. electronic bandstructure) to
several microns (e.g. internal stress)
Bulk(a thick layer of film) : the material properties will be the same as bulk properties.
Because thin films are so thin and it’s not easy to handle and that the reason we need to put them on a substrate. Most often, we will put them on a silicon wafer and that’s because it’s cheap, easy to handle and the best understood material. We can also put them on something else such as III-V material (InP, InAS, GeP, GeAs…) can also put them on metals, glass, polymer or ceramics.

## Slide 8
First atoms arrive
Nano-islands form
Finally get full layers with some roughness
Introduction – Thin film formation

## Slide 9
Introduction - Thin film applications
Thin films play a role in almost every micro- and nanostructure
Thin film functions:
Complex structures
  Micro-Electro Mechanical systems (MEMS, for example microphones and gyroscopes)
Optical
  Transparent conductive windows for solar cells
  Waveguides
Electrical
  Thin film transistors
  Piezoelectric layers
  Metallic contacts
Magnetic
Chemical
  Dopant sources
  Adhesion and surface bonding
Notes: Hvad er en MOS?
Måske deles op i to slides

## Slide 10
Introduction - Thin Film responsible persons
Henrik Nielsen
hn@dtu.dk
Roy Cork
rc@dtu.dk
Maintenance responsible persons
Peter Windmann
pewin@dtu.dk
Rebecca Ettlinger
reet@dtu.dk
Maternity leave
Pernille Larsen
pevo@dtu.dk
Patama Pholprasit
paphol@dtu.dk
Process responsible persons
Berit Herstrøm
bghe@dtu.dk
Evgeniy Shkondin
eves@dtu.dk
Contact emails:
Thin film group: thinfilm@nanolab.dtu.dk
Training: training@nanolab.dtu.dk
Darmin Catak
darcat@dtu.dk

## Slide 11
  Introduction to thin films
  Before you start – Material properties to consider
  Overview of techniques
  Equipment overview in LabAdviser
  Thin film techniques – 1st part
    Thermal oxidation
    Electroplating
Thin film TPT course – 1st part
Notes: SEM course contents

## Slide 12
Material properties that will be explained in the course:
Step coverage
Crystal structure
Film stress
Electrical properties
Composition
Material properties – Crystal structure

## Slide 13: Step coverage: What you need depends on subsequent application/processing
Line-of-sight deposition on a structured mask:
No step coverage may be wanted to allow mask dissolution (lift-off)
Si
Mask (resist)
Deposited material
Good step coverage may be desired on deeply etched structures
Some methods are better than others for uniform step coverage
Material properties – Step coverage
Conformal deposition on etched structures:

## Slide 14
200 nm
Si-wafer
resist
Ti/Au film
Ti/Au film
Cross section of high aspect ratio Si trenches covered with Al2O3 by
atomic layer deposition (ALD).
Good step coverage – conformal deposition.
Cross section of Si wafer with resist pattern + Ti/Au thin film deposited by e-beam evaporation (PVD).
No side-wall deposition (line-of-sight).
Solvent removes resist and leaves part of the Si wafer bare.
Ti/Au film
Material properties – Step coverage
Width
Depth
Aspect ratio:
Depth to width relationship of holes or trenches
Notes: Tydeliggør at kan være behov for at have en coating der dækker hele overfladen: både horisontal, veritical + hjørner
Men også kan være behov for ikke at dække sidevægge, fx ved lift-off
Evt. flere og bedre eksempler.

## Slide 15
https://en.wikipedia.org/wiki/Amorphous_silicon
Possible crystal structures:
Monocrystalline/single crystalline (e.g. Si wafers)
Continuous and unbroken crystal lattice, no grain boundaries
Poly-crystalline
Consists of multiple small crystals
Amorphous
  Non-crystalline lattice, the atoms form a continuous random network
Be aware
You may get a mix of crystal structures
Structure affected by deposition method and parameters
Post process, e.g. annealing can change the crystal structures
Material properties – Crystal structure

## Slide 16
Substrate with no film
Substrate with compressive stressed film
(film wants to expand)
Substrate with tensile stressed film
(film wants to shrink)
Possible effects of stress
Bending samples (cantilever, membranes …)
Film cracking
Film delamination
Void formation
Cracks on thick SiN film by LPCVD
Stressed cantilever
http://www.ece.eng.wayne.edu/~yxu/doc/researches/nanomechanical%20biochemical%20sensor.htm
Material properties – Film stress σ
Delamination of Ni films
Notes: Film stress:
Most films have a level of residual stress. This can be due to several reasons. It can be due to intrinsic stress like stress that evolves during film growth, for example if the surface mobility has been low during deposition then atoms can be pinned by other atoms giving intrinsic stress. It can also be caused by lattice mismatch between epitaxial layers. It can also be caused by extrinsic stress which is stress that is imposed by external factors such as temperature changes or if pressure is applied to a chip during packaging for example.
The effect of the stress is that it can bend that sample. The stress can be either compressive or tensile. A compressive film wants to expand and will make the sample convex. A tensile film wants to shrink and will make the sample concave. The result of a very stressful film can be bending of the sample which can be a problem of cantilevers and membranes, it can be film cracking, delamination of the film and void formation among other things.

## Slide 17
Doping
  Impurities provide free electrons (n-type) or ”holes” (p-type) in semiconductors
  Phosphorus for ”n” – doping and Boron for ”p” – doping.
  Not easy to measure dopant concentration directly as it is usually very low (~1013 atoms/cm3)
Conductivity/resistivity (sheet resistance, metals and semiconductors)
  Can vary across sample, thickness dependent
  Depends on doping (intended/unintended)
Insulator quality: Measure breakdown voltage
  The maximum voltage that can be applied to an insulating material before it starts to conduct
Gate oxide breakdown, Navid Azizi
Material properties – Electrical properties
Increasing voltage
No bias
Notes:   Measuring e.g. a diode can tell you about short circuits, etc.
Isulator quality :
Any insulation material remains its insulating properties up to a certain limit (There is no ideal insulator). After this limit, it starts a leakage currents go beyond limit and there is a failure insulation.
Figure: the oxide initially is almost perfect with a few in-grown traps/defects. As apply voltage generate more traps, then some percolative paths will be formed and there will be some small conductive paths through the layer. And at the end the Silicon wil melt and the Oxygen is released.

## Slide 18
  Some films have complex chemical composition
  Indium Tin Oxide (ITO)
  Al-doped Zinc Oxide (AZO)
  Silicon nitride (SixNy vs Si3N4)
Stoichiometry (proportion of components in the film)
  Affected by deposition parameters
  Can vary across substrate/depth
Surface may oxidize after deposition
Contamination
  Trace gases in vacuum systems (water, carbon compounds..)
  Contamination of source material
    e.g. many metals deposited in same chamber
  Volatile substrates/coatings from previous users
  Check cross-contamination information in LabManager!
Sputtered AZO: Al concentration (and resistivity) varies with pressure.
Crovetto et al 2016 J. Phys. D: Appl. Phys. 49 295101
N
Al
AlN with C and O incorporation
Yamamoto and Yoshida 1999, Physica B 273-274, 113-115
Material properties – Composition
Notes: Composition may not always be as you expect
-an issue with oxides, nitrides, alloys.. Any non-elemental depositions!
-sometimes it matters to your application – sometimes not
-examples are the transparent conductive oxides ITO and AZO (indium tin oxide, Al-doped Zinc oxide) as well as silicon nitride
-FIGURE Conductivity of AZO affected by Al-content, affected by deposition pressure in sputtering. Al is the open circles, look at the axis on the right of the figure.
-you can also get inhomogeneity in the composition across the substrate width due to geometry of the deposition system, or in the depth of the film if there is change over time as the film grows. Both things may happen in sputtering, but can also occur with other techniques
Surfaces (e.g. Ti, Al) can oxidize after deposition – be aware for example if you are doing characterization by optical methods
You can also get oxygen and other things you don’t want in your film during deposition
-water likes to adhere to the walls of the deposition chamber (it is adsorbed to the surfaces) and is hard to remove completely  oxidation
FIGURE example of O and C present in AlN lattice – this can change the material from n- to p-type. May or may not be intentional 
-in some systems like the sputter system, there are more than one metal source present at a time and we cannot promise that they do not contaminate each other a little bit –
-there may also (rarely) be contamination due to previous users putting volatile substances in the system that they really should not do (can happen for instance if a resist was not properly baked)
For such things, check the cross-contamination sheets found in Labmanager to make sure there will not be a problem for your application

## Slide 19
  Introduction to thin films
  Before you start – Material properties to consider
  Overview of techniques
  Equipment overview in LabAdviser
  Thin film techniques – 1st part
    Thermal oxidation
    Electroplating
Thin film TPT course – 1st part
Notes: SEM course contents

## Slide 20
Overview of techniques - How to make a thin film
Growth
Adding gas + Surface atoms  Film
Consumes the substrate
  Oxygen + Silicon wafer  Silicon oxide film
  (Dry thermal oxidation)
http://www.iue.tuwien.ac.at/phd/filipovic/node27.html
SiH4
H2
LPCVD deposition of Poly-Si
Deposition
Deposition on the surface without reaction with the substrate.
(Polycrystalline silicon, Poly-Si)
Thin films can be made either by growth or deposition
Si
Notes: Thin film can be made either by growth or deposition.
By growth, the film will be made from an adding gas plus surface atoms and that will consumes the substrate.
The most common film that made by growth is a Silicon dioxcide film and it can be made by using silicon wafer as a substrate and adding oxygen gas in a hot environment. The oxygen atom will react with silicon atom and form SiO2 .
By deposition, the film will be made by just depositing on a substrate without any reaction with the substrate.

## Slide 21
Overview of techniques - Thin film growth and deposition techniques at DTU Nanolab
Thin film deposition
Physical processes
Chemical processes
Evaporation
Sputter deposition
DC Magnetron
RF Magnetron
E-beam
Thermal
Electroplating
Chemical Vapor Deposition (CVD)
Low Pressure CVD (LPCVD)
Plasma Enhanced CVD (PECVD)
Growth
Thermal oxidation
Reactive
Ion beam
Others (diamond CVD)
Thermal
Plasma
Atomic Layer Deposition (ALD)
1st part of the course (today)
2nd part of the course
Notes: Only techniques at Danchip

## Slide 22
  Introduction to thin films
  Before you start – Material properties to consider
  Overview of techniques
  Equipment overview in LabAdviser
  Thin film techniques – 1st part
    Thermal oxidation
    Electroplating
Thin film TPT course – 1st part
Notes: SEM course contents

## Slide 23
Introduction: Thin film information in LabAdviser

## Slide 24
Introduction: Thin film information in LabAdviser
labadviser.nanolab.dtu.dk

## Slide 25
  Introduction to thin films
  Before you start – Material properties to consider
  Overview of techniques
  Equipment overview in LabAdviser
  Thin film techniques – 1st part
    Thermal oxidation
    Electroplating
Thin film TPT course – 1st part
Notes: SEM course contents

## Slide 26
Silicon dioxide (SiO2) can be grown on silicon samples in a thermal oxidation furnace.
Two types of thermal oxidation:
Dry oxidation – Oxygen reacts with Si to form SiO2:
Wet oxidation – Water vapour reacts with Si to form SiO2 :
More details will come later…
A-, B- and C-stack furnaces in the cleanroom
The A- and C- stack furnaces are used for thermal oxidation
Si(s) + O2(g) → SiO2(s)
Si(s) + 2H2O(g) → SiO2(s) + 2H2(g)
Thermal oxidation of silicon

## Slide 27
Applications - What is thermal oxide used for?
Electrical insulating layer
Protecting silicon surfaces (passivation layer)
Gate oxide layers (< 100 nm dry oxide). Essential component in the gate of a MOSFET
Optical layers
Sacrificial layer - Removed/sacrificed later in a process
  SiO2 is etched very fast in HF (hydrofluoric acid) compared to Si
Reflective SiO2 coating on solar cell
MOSFET (Metal-Oxide Field-Effect semiconductor transistor)
https://www.doitpoms.ac.uk/tlplib/semiconductors/mosfet.php
Thermal oxidation - Applications
Notes: Gate oxide: Very thin thermal oxide located over the gate or active region of individual resistors

## Slide 28
Door
Heaters
For heating zone 1,2,3
To exhaust
Scavenger box
Quartz tube
Quartz boat
Wafers
Cantilever
Gas inlet
O2
(for dry oxidation)
H2O vapour
(for wet oxidation)
N2
(for purging and annealing)
Torch/steamer/bubbler
Generates water vapour
Thermal oxidation – Furnace overview
Service area
Cleanroom

## Slide 29
http://www.processpecialties.com/thermox.htm
Furnace tube
How to generate water vapour for wet oxidation
Torch
Steamer
Bubbler
All three solutions (torch/steamer/bubbler) are being used on the oxidation furnaces in the cleanroom.
Steamer
H2 and O2 are combined in a quartz torch.
The torch is heated to > 800˚C. At this temperature H2 will ignite and burn with O2 into steam (2 H2 + 1 O2 → 2 H2O).
High quality and uniform steam
High temperature, H2 is explosive
A heater generates steam from DI water.
The steam is purified by a non-porous membrane, that only allows water molecules to pass.
Low temperature, only DI water needed
O2/N2 carrier gas is bubbled through a glass bubbler that contains DI water near the boiling point. Thus water vapour is added to the carrier gas.
Towards the gas inlet on the furnace, the water vapour is heated further to obtain steam.
Cheap and simple solution
Less clean and uniform steam
The bubbler has to be filled with water manually
Thermal oxidation – Furnace (wet oxidation)
Notes: How to generate water vapour for wet oxidation
Torch: H2 and O2 are combined in a heated to > 800˚C quartz torch. H2 ignites and burn swith O2 into steam (2 H2 + 1 O2 → 2 H2 O)
Steamer: Steam is generated from DI water by a heater. The steam is purified by a membrane, that only allows water molecules to pass
Bubbler: N2 or O2 is carrier gas is bubbled through a bubbler with DI water near the boiling point. The water vapor is heated further to obtain steam
The RHS adds controlled amounts of pure water vapor directly into any carrier gas stream. Needing only filtered house de-ionized (DI) water and power, the system delivers precise amounts of water vapor into atmospheric processes. The RHS is an integrated water vapor delivery system featuring adaptive closed loop control. Process parameters are set through automated process recipes or the controller touchscreen. Internal pressure, flow and temperature sensors provide feedback to the control system, which automatically adjusts operations to maintain required temperature and concentration stability. The RHS consists of a non-porous membrane that excludes particles, micro-droplets, volatile gases, and other opposite charged species from being transferred to the carrier gas and ensures only water vapor is added. The membrane selects only the source gas molecules. Other contaminants in the liquid source cannot permeate across the membrane or enter the carrier gas stream, resulting in a saturated product that is consistent and pure.

## Slide 30
Oxidation process:
Load Si wafers in the furnace at 700˚C (standby temperature)
Raise the temperature to 800-1150˚C (oxidation temperature)
Flow O2 or H2O vapour through the furnace to oxidize the Si wafers
Flow N2 through the furnace (for annealing and purging)
Cool down the furnace to 700˚C
Unload the wafers
Thermal oxidation - Process

## Slide 31
Recap from earlier
Two types of thermal oxidation:
Dry oxidation – Oxygen reacts with Si to form SiO2:
Wet oxidation – Water vapour reacts with Si to form SiO2 :
Process requirements:
Gas flow(s):
  O2 or H2O vapour flow
  Typical gas flow: 1-5 SLM
High temperature:
  Typical temperature: 800-1150˚C
Atmospheric pressure:
  ~1 bar = 100 hPa (vacuum not required)
Thermal oxidation – Process requirements
Si(s) + O2(g) → SiO2(s)
Si(s) + 2H2O(g) → SiO2(s) + 2H2(g)
DTU Nanolab has fixed recipes on most oxidation furnaces, e.g.
”WET1000” – Wet oxidation at 1000˚C
”DRY1100” – Dry oxidation at 1100˚C
Users can change the oxidation time in the recipes to decide how thick an oxide layer they want to grow

## Slide 32
Thermal oxidation – Properties
Process properties:
Growth on both sides of the wafers
Batch process
Good uniformity (<1.5 % non-uniformity over 30 4” wafers)
Very reproducible
General thermal oxide properties:
Very good electric insulator
Can be doped with boron or phosphorous
(using PECVD, ion implantation or pre-deposition in other furnaces at DTU Nanolab)

## Slide 33
Thermal oxidation – Advantages and limitations
Dry oxidation
Advantages:
Very dense amorphous film
High quality film with few defects
Limitations:
Slow oxidation rate
Maximum oxide thickness at DTU Nanolab ~300 nm
Wet oxidation
Advantages:
Faster oxidation rate
Maximum oxide thickness at DTU Nanolab ~3 µm
Limitations:
Less dense amorphous film
Why does dry oxide have a better quality than wet oxide?
H2 causes imperfections in wet oxide layers, and therefore these are less dense than dry oxide layers
Si(s) + O2(g) → SiO2(s)
Si(s) + 2H2O(g) → SiO2(s) + 2H2(g)
Notes: In wet oxides, hydrogen atoms liberated by the decomposition of the water molecules during oxidation, produce imperfections that degrade the oxide quality somewhat

## Slide 34
Important:
SiO2 is grown (not deposited like e.g. PVD, CVD and ALD layers) on Si wafers
The O2 or H2O vapour molecules react with Si at the interface between the oxide layer and the Si surface. I.e. the oxidants have to diffuse through the already grown SiO2 layer to reach this interface
Advantage:
Contaminations on the Si surface moves to the outer SiO2 surface, thus the grown SiO2 layer is very clean
http://www.processpecialties.com/thermox.htm
Oxidation reaction
SiO2 layer
Thermal oxidation - Process

## Slide 35
Question
Do you think that the growth rate for SiO2 is linear as function of time? – Why or why not?
There is no quick answer to this question…
The growth rate for very thin (some hundred nm) oxide layers is linear.
For thicker oxide layers the growth rate is not linear and decreases with time:
When the oxide layer grows, it takes longer and longer time for the O2 or H2O vapour molecules to diffuse through the already grown oxide layer and reach Si surface to form more SiO2
Thermal oxidation - Process
http://www.processpecialties.com/thermox.htm
Oxidation
reaction
SiO2 layer

## Slide 36
BREAK

## Slide 37
The oxidation reaction occurs a the interface between the oxide layer and the Si surface
To understand the oxidation process, we first have to understand the kinetics of the oxide growth, i.e. how the oxidants are being transported to the Si surface.
Oxidants: O2 or H2O vapour molecules
What is going to happen for the oxidation to the place?
Oxidants in the gas steam are transferred to the gas-SiO2 interface
The oxidants diffuse through the already grown oxide layer to reach the SiO2-Si interface
The oxidants reach with Si at the SiO2-Si interface
Thermal oxidation – Process (Deal-Grove model)
http://www.processpecialties.com/thermox.htm
SiO2-Si interface
Gas-SiO2 interface
SiO2 layer
Gas steam

## Slide 38
1.
The flux of oxidants in the gas steam is governed by concentration differences:
Or
2.
The flux is related the diffusion of oxidants through the existing oxide layer.
The oxidants moves because of a concentration gradient:
3.
The flux is related to the reaction between the Si and oxidants at the Si-SiO2 interface.
The reaction is governed by the reaction rate of the oxidants:
Or
,
SiO2 layer
Si
Gas steam (O2/H2O)
CG
Ci
CO
Oxide thickness x
CS
Concentrations of oxidants (number of molecules per area):
in the bulk of the gas
right next to the gas-oxide interface
at the outer oxide surface
at the inner oxide surface
Thermal oxidation – Process (Deal-Grove model)
N1 values:
Dry oxidation: N1 for O2 = 2.2·1022
Wet oxidation: N1 for H2O = 4.4·1022
F2
F3
F1
Three fluxes are related to the oxidation process:
Flux: Number of molecules per time per area
Notes: Flux: Number of molecules per second per cm2 (1/(cm2*s))
The fluxes are related to the concentration of the oxidants in the various regions:
C_G: Concentration of oxidants in the bulk of the gas
C_S: Concentration of oxidants right next to the gas-oxide interface
C_O: Contration of oxidants at the outher oxide surface
C_i: Contration of oxidants at the inner oxide surface
C*: Eqiulibrium concentration of the oxidizing species in the oxide
1.
h_G: Proportional constant/gas phase mass transfer coefficient
Ideal gas law: F1=(h_G/(kT))*(P_G-P_S)
Henri’s law: In equilibrium, the concentration of species within a solid is proportional to the partial pressure of that species in the surrounding gas)
3.
K_S: Proprotional constant/reaction rate constant

## Slide 39
Steady state condition:
From this equation the oxide growth characteristic, i.e. how the SiO2 thickness increases with time, can be found.
The slowest factor (F1, F2 or F3) now becomes the limiting factor that determines the overall growth rate.
F1 is not a limiting factor.
With some math, the equation can be written as:
Thermal oxidation – Process (Deal-Grove model)
This equation is called the linear-parabolic Deal-Grove equation, and it predicts the oxide growth very precisely
(represents the time that would have been needed to grow the oxide layer
that was already present on the Si surface when the oxidation begins)
t: Oxidation time
xi: Initial oxide thickness (oxide present when the oxidation begins)
Notes: Three are three rates to consider:
The rate at which the gas molecules arrive at surface
The rate at which molecules diffuse through the already formed oxide layer
The rate at which the reaction occurs at the interface
In steady-state, all three rates are the same. Then the slowest rate becomes the limiting factor that determines the overall growth rate
Not a limiting factor
Limiting factor in the parabolic growth regime
Limiting factor in the linear growth regime

## Slide 40
Deal-Grove equation:
The oxide thickness is found by finding the roots (for which x > 0) of the equation:
Short oxidation time
x is linear proportional to time, and thus the oxidation rate follows a linear growth rate curve.
The oxidation rate in limited by how fast the oxidants can react with the Si atoms at the SiO2-Si interface. I.e. F3 is the limiting factor.
https://www.semitracks.com/newsletters/october/2014-october-newsletter.php
Long oxidation time
The oxidation rate follows a parabolic growth rate curve.
The oxidation rate is limited by diffusion, because it takes longer and longer times for the oxidant to diffuse through the oxide layer. I.e. F2 is the limiting factor.
Thermal oxidation – Oxide growth rate (Deal-Grove model)
Notes: The Deal-Grove model matematically describes the growth of an oxide layer on a Si substrate
The model assumes that the oxidation reaction occurs a the interface between the oxide layer and the Si surface

## Slide 41
Oxide thickness as function of oxidation time for different oxidation temperatures
The graphs show experimental data form the A1 Boron Drive-in and Pre-dep furnace in the cleanroom
The curves are fitted from the numbers in the process log
Oxidation time [min]
Oxide thickness [nm]
Oxide thickness [nm]
Oxidation time [min]
Wet oxidation
Si (100) wafers
Dry oxidation
Si (100) wafers
Thermal oxidation – Experimental data

## Slide 42
From the graphs it is seen that:
Bothand B increase with temperature
  So, increase temperature > Faster oxidation rate > Thicker oxide layer
Both and B are larger for wet oxidation than for dry oxidation. Thus, wet oxide is grown faster than dry oxide.
, and C* is much larger for wet oxidation than for dry oxidation,
And H2O molecules form hydroxide (OH-), that diffuses much faster through the already-grown oxide layer than O2
Linear growth regime:
Diffusion-limited regime:
Thermal oxidation – Oxide growth rate (Deal-Grove model)
C* values (for a given T): Dry oxidation: C* = 5.2·1016
Wet oxidation: C* = 3.0·1019
https://nptel.ac.in/courses/103106075/30
B (µm2/hr)
B/A (µm/hr)

## Slide 43
https://www.tf.unikiel.de/matwis/amat/elmat_en/kap_6/backbone/r6_2_1.html
Experimental data points
Extrapolated linear curve fitted to the Deal-Grove model
(assumption: xi = 21 nm at t = 0)
Linear growth regime
Thermal oxidation – Limitations of the Deal-Grove model
Limitations of the Deal-Grove model:
Very thin (<30 nm) dry oxide layers are grown faster than the
model predicts, and they grow non-linearly.
The model is only valid for single crystal silicon
  Poly-crystalline Si is oxidized faster than the model predicts, because oxidants diffuse rapidly along grain boundaries. Furthermore, the random orientation of the crystal grains makes it difficult to chose a value for the linear rate constant
Heavily doped Si wafers might be oxidizes faster than the model predicts
  More info will come later…

## Slide 44
Question:
How can you tell how thick a oxide layer is?
Measure the thickness
  Use an ellipsometer or spectrometer (Filmtek) or use an SEM (cleave the wafer first)
Look at the wafer
  The colour is determined by the SiO2 thickness
  The eye can tell the colour difference between two SiO2 layers having a 10 nm thickness difference
http://scme-nm.org/files/Rainbow_Activity_PG.pdf
Si wafer with different SiO2 thicknesses
Thermal oxidation – How find the oxide thickness

## Slide 45
Thermal oxide has an amorphous/porous crystal structure
43% of the space is occupied by SiO2 molecules
However, close to the Si/SiO2 interface thermal oxide has a crystalline structure.
https://www.nde-ed.org/EducationResources/CommunityCollege/Materials/Structure/solidstate.htm
Quartz crystal lattice
Density 2.65 g/cm3
Thermal SiO2
amorphous structure
Density 2.21 g/cm3
Thermal oxidation – Crystal structure
Si
Thermal SiO2
TEM image of Si/SiO2 interface
https://www.slideshare.net/bhargavveepuri/61-thermal-oxidation-12micro-tech2013
Notes: Silicon dioxide has a tetrahedal structure

## Slide 46
The amorphous crystal structure of the SiO2 layer allows O2 and H2O molecules to diffuse through it.
Also impurities, e.g. sodium and potassium, can diffuse through the SiO2 layer.
So even a very small amount of impurities will affect the cleanliness of a furnace.
How to estimate the oxide quality?
Measure the break-down voltage (easy)
Measure the life-time (best, but time consuming as dedicated devices have to be made)
Thermal oxidation – Oxide quality
It is only possible to grow a high quality oxide with a low impurity level in a very clean furnace
Before you use a furnace
Check the cross contamination information in LabManager
Check the RCA cleaning rules
And ask the responsible persons, if you have questions
Example
The A-stack furnaces are some of the cleanest machines in the cleanroom.
Only Si wafers are allowed in these furnaces, and all wafers have to be RCA cleaned.

## Slide 47
O2 or H2O vapour
The SiO2 growth consumes the Si surface, and thus expands the volume
Example:
How much Si is consumed to grow 100 nm of SiO2?
I.e. 45% of the oxide thickness is below the original Si surface,
and 55 % is above it
How much is the volume expanding?
Thermal oxidation – Si consumption and volume expansion

## Slide 48
  https://www.kth.se/social/upload/4f283b68f276541a2900000e/Lecture%205%20-%20Spring%202012.pdf
Before oxidation
After oxidation
When a Si wafer is being oxidized, the volume can only expand upwards (not to the sides)
Therefore thermal oxide has a high compressive stress
Thermal oxidation – Stress
Compressive stress
– the SiO2 layer wants to expand
However, the stress will cancel out when oxide is grown on both sides of a wafer

## Slide 49
Thermal oxidation – Oxidation of structured samples
Cross section of 700 nm – 2 µm wide trenches with scallops
Structured Si sample that has been oxidized
For structured samples, the oxide thickness will be smaller in edges compared to flat areas
This is an advantage for e.g. for removal of scallops after DRIE (deep reactive ions etching)
But be aware of that Si is being consumed during the oxidation, and very narrow structures might therefore be completely oxidized
Scallops after
oxidation
Scallops before
oxidation
Cross section of 700 nm – 2 µm wide trenches with scallops
1500 nm
Structured Si sample that has been oxidized
Notes: Færdig

## Slide 50
Question:
Why is the oxide layer thicker on the sidewalls than in the bottom of the structures?
The oxidation rate is not same for all crystal directions
1150 nm
995 nm
Structured Si sample that has been oxidized
Thermal oxidation – Oxidation of structured samples
Si (100)
Si (110)

## Slide 51
The oxidation rate depends the crystal orientation, i.e. on the amount of available Si atoms that the O2/H2O vapour molecules can react with:
Si (111) has more available atoms per unit area on the surface than Si (100).
Therefore, Si (111) oxidizes faster than Si (100) is the linear rate regime, where the oxidation rate is limited by the reaction rate.
Reaction rate for the linear rate constant in the Deal-Grove model:
  In the diffusion-limited regime, this advantage is somewhat lost, because the oxidation rate is limited by diffusion
  At DTU Nanolab mostly (100) wafers are being oxidized
http://www.iue.tuwien.ac.at/phd/hollauer/node14.html
Thermal oxidation – Growth rate as function of crystal structure
Notes: In parabolic rate regime, the curves come closer together, because the oxidation rate is diffusion dependent
of silicon atoms available for the reaction. It was found experimentally that surfaces which provide more available reaction sites to silicon have a higher oxidation rate [122]. The ratio for the parabolic rate constant in silicon crystal orientations (111):(100) was found to be 1.68:1.
Similarly, it has been suggested that the ratio for the parabolic rate constant in crystal orientations (110):(100) is approximately 1.45:1, noticeable on thicker oxide along sidewall surfaces.

## Slide 52
For heavily doped wafers, the oxidation rate is faster than the Deal-Grove model predicts
Boron doped wafers (p+ doping):
Faster oxidation rate
  Boron converts the oxide into borosilicate glass (boron-doped SiO2) that strains the lattice constant,
  so the O2 /H2O molecules diffuse faster through the layer than through undoped SiO2
Phosphorous doped wafers (n+ doping):
Faster oxidation rate, but only in the reaction limited regime
  A defect region is created at the Si-SiO2 interface, and this increases the reaction rate
Thermal oxidation – Oxidation of doped Si wafers

## Slide 53
An annealing is normally done after the oxidation
Advantages of the annealing:
Improve the oxide density
Repair atomic defects within the crystal structure
Remove dangling bonds at the Si-SiO2 interface
Reduce vacancies and dislocations at grain boundaries
At DTU Nanolab the annealing time is normally 20 minutes, but the time can be adjusted.
The annealing is done at the same temperature as the oxidation.
Thermal oxidation – Annealing after oxidation
Notes: Virker annealing med N2? H2/forming gas bedst
Advantages of the annealing:
Improved the oxide density
Repair atomic defects within the crystal structure (from the Si wafer microsctructure?)
Remove dangling bonds at the Si-SiO2 interface
Reduce vacancies and dislocations at grain boundaries

## Slide 54
A-stack furnaces:
A1 Boron Drive-in and Pre-dep Furnace
  Dry and wet dry oxidation (using a torch) of p-type wafers
  Sample size: 4” wafers
A2 Gate Oxide Furnace
  Dry oxidation of very clean wafers
  Sample size: 4” wafers
A3 Phosphorous Drive-in Furnace
  Dry and wet oxidation (using a torch) of n-type wafers
  Sample size: 4” wafers
A-stack furnaces
Thermal oxidation furnaces at DTU Nanolab

## Slide 55
C-stack furnaces:
C1 Anneal Oxide furnace
  Dry and wet oxidation (using a steamer)
  Sample size: 4” and 6” wafers
C2 III-V Oxidantion furnace
  Wet oxidation (using a bubbler) of samples with III-V materials
  Sample size: 4” and small samples
  Being tested at the moment
C3 Anneal Bond furnace:
  Dry and wet oxidation (using a bubbler)
  Sample size: 4” wafers
C4 Al-Anneal furnace
  Dry oxidation of wafers with Al, Al2O3 and TiO2
  Sample size: 4” wafers
C-stack furnaces
Thermal oxidation furnaces at DTU Nanolab

## Slide 56
Multipurpose Anneal furnace
Dry oxidation of ”dirty” samples (e.g. samples with metals)
Sample size: Small samples, 4” and 6” wafers
Noble furnace
Dry oxidation (using a bubbler) of ”dirty” samples, e.g. samples with metal
Sample size: Small samples, 4” wafers
Multipurpose Anneal furnace
Noble furnace
Thermal oxidation furnaces at DTU Nanolab
Furnace reponsible persons

## Slide 57
  Introduction to thin films
  Before you start – Material properties to consider
  Overview of techniques
  Equipment overview in LabAdviser
  Thin film techniques – 1st part
    Thermal oxidation
    Electroplating
Thin film TPT course – 1st part
Notes: SEM course contents

## Slide 58
Electroplating
At DTU Nanolab, thick Ni layers (up to 700 µm) is deposition by use of electroplating
Ni electroplating is typically used to transfer structures from Si wafers to Ni shims
Application examples:
Stamps for injection moulding (≈350 µm Ni)
Foils for embossing/imprinting (100-200 µm Ni)
https://www.researchgate.net/figure/Replication-techniques-From-the-lithographically-prepared-master-it-is-possible-to-make_fig9_221909895
Process:
Make structured Si wafer (with seed layer)
Ni electroplating
Remove Si wafer to release Ni shim
Electroplating is an electrochemical deposition technique for deposition of thick metal layers
Ni
Si

## Slide 59
Electroplating – Process (electroplating of Ni)
A sample is lowered into an electrolyte (an aqueous solution) containing Ni ions.
The electrolyte has a well-controlled temperature and pH that are adjusted automatically:
Temperature: 52˚C
pH: ~3.7 (acidic)
A voltage is applied between a titanium basket with Ni pellets and the sample, i.e. across an anode and a cathode
At the anode, metallic Ni is oxidized to Ni ions:
At the cathode, Ni ions from the electrolyte is reduced to metallic Ni:
Anode (+)
Titanium basket with Ni pellet
Cathode (-)
Sample
Electrolyte containing Ni ions
Electroplated Ni layer
Conducting seed layer
Si wafer
Electroplating is an electrochemical deposition technique
Notes: The plating bath is an aqueous solution of nickel sulfamate, borid acid and sulfamate acid
Ni ions are transported from anode to cathode
Ni is deposited at the cathode, when positively changed Ni ions are discharged

## Slide 60
Electroplating – Sample requirements
Sample requirements:
Sample size: 50, 100 and 150 mm wafers
Wafer thickness: ~350-1000 µm
Sample materials must be compatible with the electrolyte
Electroplating can only be done on wafers with a conducting surface
For instance, Si wafers can be covered by a thin (~50-100 nm) seed layer of Ti/Cr+Au or NiV (nickel-vanadium) using PVD deposition.
Important:
All parts of the wafer surface, including sidewalls on high aspect ratio structures, must covered with a uniform conducting seed layer.
Thus, sputtering is the preferred PVD method to cover wafers with a conducting seed layer, because it gives a good step coverage.
More info in the 2nd part of this course…
A dedicated wafer holder has to be used in the sputter coater to ensure deposition in the edge of the wafer surface - and thus electric edge contact in the electroplater.
Electroplated Ni layer
Conducting seed layer
Si wafer
Electroplater sample holder with 4” wafer
Notes: Nickel-Vanadium is used, becaues nickel is oxidizing too fast and becomes non-conductive
Au alone on Si is not allowed (poor adhesion, delaminates and causes clogging of filters).

## Slide 61
Electroplating – Process
The deposition rate is controlled by:
The charge [Ah]
The current density [A/dm2], i.e. current and the wafer size
The deposition rate increases with the current density
But a low current density often has to be used - Why?
Improve filling of structures
Better uniformity
Lower surface roughness
More hard surface

## Slide 62
Time [minutes]
Current [A]
Electroplating – Process example
Example:
DTU Nanolab has three standard programs for deposition of 350 µm Ni on 4” Si wafers with ~100 nm NiV.
The maximum current [Imax] and thus the deposition
time is different for the three programs
The charge is 18.3 Ah for all the programs
The depositions starts with a slow current ramping to avoid wafer (and sample holder) damage, as the seed layer cannot conduct high currents.

## Slide 63
Electroplating at DTU Nanolab
Typical process results
Thickness of deposited Ni layer: ~50-700 µm
Uniformity of deposited Ni layer: 2-10%
Contact email:
wetchemistry@nanolab.dtu.dk
Electroplater (Electroplating-Ni) in the cleanroom
Wafer mounted in a sample holder above the electrolyte
Claus, choi@dtu.dk
Ni shim with micro channels made by electroplating

## Slide 64
Thanks for your attention
Please send feedback, questions and comments to thinfilm@nanolab.dtu.dk
Or contact one of the responsible persons
For training, please contact training@nanolab.dtu.dk