## Slide 1: Test to evaluate number of particles/possible pinholes on metal films in the Temescal E-beam evaporator and the Wordentec
By Evgeniy Shkondin and Rebecca Ettlinger, December 2018
(all SEM images by Evgeniy)

## Slide 2
DF, Mag x10
DF, Mag x10
Normal, Mag x10
DF, Mag x100
Normal, Mag x100
Temescal TiAu QC (10 nm Ti + 90 nm Au deposited at 10 Å/s)
27-28.11.2018
Images from optical microscopy (Dark field and Bright field)
Bright spots seen in dark field – points that scatter light on the smooth surface. Could be pinholes or particles.

## Slide 3
Particles in a 10+90 nm TiAu QC wafer
Temescal 27-28.11.2018
SEM images of same wafer as in the previous slide to verify whether spots seen in optical microscopy are from particles or pinholes
Lots of particles seen. So majority probably particles.

## Slide 4
Particles in 50 nm Al (not as many as in TiAu case)
Temescal 26.11.2018 (after service). Standard recipe.

## Slide 5
5+45 nm TiAu
50 nm Al
50 nm Al
5+45 nm TiAu
Grains on deposited films

## Slide 6: Silver. Optical microscope DF Mag x100. Temescal 50 nm 2Å/s

## Slide 7
Particle count from KLA Tencor Surfscan using the program 4’’ test (particle cutoff size is 500 microns)
Material | Thickness | Dep. rate | Number of particles | Comments
Ag (Thermal) | 200 nm | 2 Å/s | - | (mask 8 pallets)
Al | 50 nm | 2 Å/s | 19 | 
TiAu | 5+45 nm | 2 Å/s | 1820 | 
Ni | 50 nm | 2 Å/s | 22 | 
Wordentec (23.11.2018)
Material | Thickness | Dep. rate | Number of particles | Comments
Al | 100 nm | 2 Å/s | 135 | 
Blank wafer |  |  | 10 | Extra movement of wafer and wafer holder before load
Blank wafer |  |  | 6 | 
Wordentec (6.12.2018)

## Slide 8
Material | Thickness | Dep. rate | Number of particles | Comments
Al | 50 nm | 2 Å/s | 416, 670, 1176* | * Vaccum tweezer fingerprint
TiAu | 5+45 nm | 2 Å/s | 930, 1146, 993 | 
Temescal 20.11.2018
Temescal 21.11.2018
Material | Thickness | Dep. rate | Number of particles | Comments
Al | 50 nm | 2 Å/s | 104, 208, 54 | Measured several days later
Ag | 50 nm | 2 Å/s | 4740, 3602, 3787 | Measured several days later. Optical microscope reveals big grains

## Slide 9
1) Load-pump-gate open/close-vent-unload | 93, 157
2) Load-pump-vent-unload, no rotation | 51, 94
3) Load-pump-vent-unload, with rotation | 33, 35
Temescal 26-27.11.2018 (after service)
Material | Thickness | Dep. rate | Number of particles | Comments
4) Al | 50 nm | 2 Å/s | 109, 210 | 
5) Al | 50 nm | 2 Å/s | 920, 2485 | Longer soak-ramp time i)
6) Al | 50 | 2 Å/s | 63, 204 | Short soak-ramp time ii)
7) TiAu | 5+45 | 2 Å/s | 397, 661* | * A bit of a mess during the loading
Standard |  | (i) | Used instead | 
R1 | 30 s |  | R1 | 60 s
S1 | 20% |  | S1 | 18%
S1 time | 30 s |  | S1 time | 60 s
 |  | ---- |  | 
R2 | 30 s |  | R2 | 60 s
S2 | 33% |  | S2 | 33%
S2 time | 120 s |  | S2 time | 240 s
Standard |  | (ii) | Used instead | 
R1 | 30 s |  | R1 | 30 s
S1 | 20% |  | S1 | 18%
S1 time | 30 s |  | S1 time | 30 s
 |  | ---- |  | 
R2 | 30 s |  | R2 | 30 s
S2 | 33% |  | S2 | 33%
S2 time | 120 s |  | S2 time | 30 s

## Slide 10
Temescal 27-28.11.2018
Material | Thickness | Dep. rate | Number of particles | Comments
TiAu | 10+90 nm | 10 Å/s | 6895 | QC
I) Load-pump-gate open-shutter open/close3x –gate close-vent-unload | 32, 135
Material | Thickness | Dep. rate | Number of particles | Comments
II) TiAu | 10+90 nm | 10 Å/s | 1575, 1606 | Longer soak-ramp time iii)
III) Ti | 10 nm | 10 Å/s | 599, 1628 | Parameters from recipe in (II)
IV) TiAu | 10+90 nm | 2 Å/s | 789, 816, 925* | The same as (II) but with lower dep. rate *Not proper fixed in the holder
V) Ti | 10 nm | 2 Å/s | 53, 70 | The same as (III) but with lower dep. rate
Standard |  | (iii) | Used instead | 
R1 | 30 s |  | R1 | 60 s
S1 | 20% |  | S1 | 20%
S1 time | 30 s |  | S1 time | 60 s
 |  | ---- |  | 
R2 | 30 s |  | R2 | 60 s
S2 | 31% |  | S2 | 33%
S2 time | 60 s |  | S2 time | 60 s

## Slide 11
Temescal 30.11.2018
Test on wafers from load-pump-gate open-gate close-vent-unload with clean sample holders | Approx. 350
Temescal 13.12.2018
Test on wafers from load-pump-gate open-gate close-vent-unload | 64, 71, 104
Vacuum cleaner applied for the load lock area. Test on wafers from load-pump-gate open-gate close-vent-unload | 71, 90, 101

## Slide 12: Observations
Load / unload in Temescal gives more particles than load/unload in Wordentec
Single Ti/Au film in Wordentec had more particles than TiAu films from Temescal
Single Al films from Wordentec had lower numbers of particles than films from Temescal (but comparable when Temescal soak time for Al was reduced).
Ti/Au films have many more particles (order of magnitude more) than Ni or Al
Ag was not measured well- large number of particles may have been due to large grains especially after oxidation
With Al, we got more particles when we doubled the soak and ramp times compared to normal
With TiAu, we got fewer particles from doubling the soak and ramp times for Au.
With TiAu in Temescal and Al in Wordentec we got fewer particles from reducing the deposition rate from 10 Å/s to 2 Å/s
”Identical” Al deposition particle count varied hugely – from 54 to 670 !! – so also just lots of random variation…
After cleaning gate valve (service with Hugh) the number of particles in Al deposition was similar to before but the number of particles in TiAu appeared to decrease…
Load/unload in Temescal: similar number of particles whether or not rotation was on and whether or not gate valve was opened/closed. And whether or not shutter to source was opened closed. Similar number of particles after Al deposition under best circumstances as from just load/unload.

## Slide 13: Conclusions
Lower deposition rate (2 Å/s) gives fewer particles than high (10 Å/s)
Ti/Au films have many more particles (order of magnitude more) than Ni or Al
Optimized soak and ramp times can decrease particle number
Temescal starting point is more dirty than Wordentec – with optimization probably all processes in Wordentec would be cleaner than in Temescal
For Temescal: Cleaning the gate valve did not seem to help but perhaps cleaning heath or something else near source improved particle count for TiAu after service
For Temescal: Operating gate valve, hula rotation and source shutter does not add significantly to number of particles (maybe some increase from gate valve). Even depositing Al with best recipe did not add significantly to the number of particles compared to simple load/unload.
-- so more particles in Temescal than Wordentec simply come from loading/unloading/pumping down/venting
This may have to do with the fast, turbulent venting of the Temescal.

## Slide 14: Outlook
Need to see if results are stable over time (does it depend a lot on previous use?)
Can test if vacuum cleaning before loading helps
Could geometry of venting of loadlock be reason for more particles in Temescal than Wordentec?
Could using dummy wafers help prevent particles released from hula? (parts of hula above center of samples is hard to clean and it looks like there might be tiny flakes released from here)
Can ask users to put in old CVD qc wafers with oxide as dummies 
In March 2019 there will be major service – test particles again before and after the service (maybe also after a few depositions after service where any particles from cleaning are encapsulated)