## calculations
Uncertainty budget for DektakXTA (last edit January 2022)
General
source of uncertainty | value +/- | shape of pdf | divisor | standard uncertainty | What is this uncertainty about?
uncertainty on calibration of standard step height | 0.072 µm for the 24.925 µm standard and 17 nm for the 917 nm standard (95 % confidence limits) | Gaussian | 2 | depends on range. Systematic error | difference between what we expect that the standard step height is and what it truly is
uncertainty on QC measurements of the standard step height. We do calibration if outside limits | Treating the distribution of actual measured step heights as Gaussian distributions gives a "standard deviation" of 25 Å for the small range and 66 nm for the 65.5 micron range (after excluding the worst outliers). | Actual measurements of the standard step height are more rectangular-distributed than Gaussian-distributed. They are centered on 916 nm for the 6.5 micron range and 24871 nm for the 65.5 micron range. Due to our QC procedure the distribution cuts off at 0.15 µm for the 24.865 µm standard (used to calibrate the 3 largest Dektak ranges) and 12 nm Å for the 917 nm standard (smallest range) | NA | can only be estimated.…could be a combination of systematic and random error | difference between the average value we measure for the standard step heigth with the Dektak and what we expect the standard step size to be
resolution | 1 Å, 10 Å, 80 Å or 160 Å (for ranges 65 kÅ, 655 kÅ, 5240 kÅ and 1 mm respectively) | uniform | sqrt 3 | depends on range. Random error | difference between the value the Dektak spits out and what it would spit out with better resolution…
standard uncertainty of mean (10 repeated readings) | 24 Å on 6.5 micron range, 45-65 nm on larger ranges for the standard samples. Could be larger for other sample types | Gaussian | none | depends on range and on sample. Random error | difference between what we measure in one Dektak scan of some feature of interest and what we measure as the average of many scans
uncertainty on reading w 6.5 micron range
source of uncertainty | value +/- (nm) | shape of pdf | divisor | standard uncertainty (nm)
standard step height uncertainty | 17 | Gaussian | 2 | 8.5
QC uncertainty | 2 | rectangular-ish distribution cut off at 12 nm; see graph on right (expected value 916 nm 2015-2020) | 1 | 2
resolution | 0.1 | uniform | sqrt 3 | 0.057735
standard uncertainty of mean | 2.4 | Gaussian | 1 | 2.4
total uncertainty |  |  |  | 9.0561
95 % confidence int |  |  |  | 18.1122
uncertainty on reading w 65.5 micron range
source of uncertainty | value +/- (micron) | shape of pdf | divisor | standard uncertainty (microns)
standard step height uncertainty | 0.072 | Gaussian | 2 | 0.036
QC uncertainty | 0.062 | 0.06 | 1 | 0.062
resolution | 0.001 | uniform | sqrt 3 | 0.00057735
standard uncertainty of mean | 0.045 | Gaussian | none | 0.045
total uncertainty |  |  |  | 0.084648
95 % confidence int |  |  |  | 0.1693
uncertainty on reading w 524 micron range
source of uncertainty | value +/- (micron) | shape of pdf | divisor | standard uncertainty (microns)
standard step height uncertainty | 0.072 | Gaussian | 2 | 0.036
QC uncertainty | 0.065 | rectangular or Gaussian-ish distribution cut off at +/-150 nm from expected value due to our QC limits; see graph on right (expected value 24865 nm 2015-2020) | 1 | 0.065
resolution | 0.008 | uniform | sqrt 3 | 0.0046188
standard uncertainty of mean | 0.045 | Gaussian | none | 0.045
total uncertainty |  |  |  | 0.08699
95 % confidence int |  |  |  | 0.17398
uncertainty on reading w 1 mm range
source of uncertainty | value +/- (micron) | shape of pdf | divisor | standard uncertainty (microns)
standard step height uncertainty | 0.072 | Gaussian | 2 | 0.036
QC uncertainty | 0.042 | rectangular or Gaussian-ish distribution cut off at +/-150 nm from expected value due to our QC limits; see graph on right (expected value 24865 nm 2015-2020) | 1 | 0.042
resolution | 0.16 | uniform | sqrt 3 | 0.092376
standard uncertainty of mean | 0.045 | Gaussian | none | 0.045
total uncertainty |  |  |  | 0.10888
95 % confidence int |  |  |  | 0.21776
This chart isn't available in your version of Excel. Editing this shape or saving this workbook into a different file format will permanently break the chart.
This chart isn't available in your version of Excel. Editing this shape or saving this workbook into a different file format will permanently break the chart.
This chart isn't available in your version of Excel. Editing this shape or saving this workbook into a different file format will permanently break the chart.
This chart isn't available in your version of Excel. Editing this shape or saving this workbook into a different file format will permanently break the chart.

## data small range
6.5 µm range
Date | User | ID | Hardware | Average (Å) | Deviation (Å)
2018-04-20 | camtin/reet |  |  | 9144 | 3.4
2018-04-23 | reet
2018-04-25 | reet |  |  | 9153 | 16.1
2018-06-01 | reet |  |  | 9155.1 | 11.9
2018-06-01 | reet
2018-06-28 | paphol |  |  | 9158.64 | 12.25
2018-07-03 | reet |  |  | 9141.4 | 45.14
2018-07-26 | reet |  |  | 9148.43 | 18.44
2018-07-30 | paphol |  |  | 9120 | 48.18
2018-08-06 | reet |  |  | 9181.19 | 11.15
2018-09-05 | reet |  |  | 9193.56 | 26
2018-10-08 | paphol |  |  | 9180.956 | 26.064
2018-11-05 | reet |  |  | 9199.7 | 25.48
2018-11-21 | reet |  |  | 9146 | 13
2018-12-11 | eves/reet |  |  | 9149.26 | 50.75
2019-01-11 | eves |  |  | 9142.6 | 16.234
2019-02-18 | paphol |  |  | 9160.36 | 15.7
2019-03-12 | eves |  |  | 9150.38 | 29.91
2019-03-13 | eves |  |  | 9150.38 | 29.91
2019-03-14 | paphol |  |  | 9177.386 | 18.591
2019-04-17 | eves |  |  | 9163.45 | 33.69
2019-05-23 | eves |  |  | 9178.891 | 14.226
2019-06-21 | eves |  |  | 9206.185 | 21.516
2019-07-25 | eves |  |  | 9158.09 | 50.2355
2019-08-26 | paphol |  |  | 9178.86 | 37
2019-10-02 | paphol |  |  | 9190.505 | 18
2019-11-05 | eves |  |  | 9154.58 | 59.26
2019-12-03 | paphol |  |  | 9156.043 | 18.31
2020-01-06 | eves |  |  | 9154.993 | 13.123
2020-02-04 | paphol |  |  | 9151.716 | 17.423
2020-02-18 | reet |  |  | 9159 | 12
2020-03-09 | reet |  |  | 9189 | 32
2020-04-23 | paphol |  |  | 9189 | 12
2020-05-25 | reet |  |  | 9191 | 15
2020-06-25 | reet |  |  | 9186 | 46
2020-07-29 | reet |  |  | 9183 | 13
2020-08-28 | paphol |  |  | 9184 | 12
2020-09-24 | reet |  |  | 9176 | 17
2020-10-07 | reet |  | new stylus | 9187 | 14
|  |  |  | 9167.288 | 23.6211
|  |  |  | 20.0541 | 13.8704

## data large ranges
data w outliers removed* |  |  | 65.5 µm range |  | 524 µm range |  | 1 mm range |  | *The outliers that are removed are
Date | User | Hardware | Average (nm) | Deviation (nm) | Average (nm) | Deviation (nm) | Average (nm) | Deviation (nm) | *Data where we had to recalibrate the dektak
2018-04-23 | camtin/reet |  | 24763.8 | 69.1 |  |  | 24757.1 | 23.2 | *Data where the standard error is too large and some part of the data obviously should have been re-measured
2018-04-23 | reet |  |  |  | 24907.6 | 20.9
2018-05-31 | reet |  | 24931.7 | 44.1 |  |  | 24840.3 | 39.9
2018-06-01 | reet
2018-06-01 | reet |  |  |  | 24864.6 | 44.8
2018-06-26 | paphol |  | 24871.4 | 58.7 | 24792.99 | 32.39 | 24797.17 | 35.26
2018-07-03 | reet |  | 24782.47 | 32.58 | 24763 | 35.47 | 24793.86 | 42.07
2018-07-26 | reet |  | 24740.9 | 36.8 | 24719.54 | 38.32 | 24756.95 | 35.76
2018-07-30 | paphol
2018-07-30 | paphol |  | 24871.04 | 33.48 | 24853.35 | 41.36 | 24762.74 | 34.71
2018-08-06 | reet | new stylus | 24961.84 | 36.08 | 24951.8 | 34.3 | 24850.46 | 36.26
2018-09-05 | reet |  | 24959.88 | 34.32 | 24955.27 | 35.44 | 24853.68 | 35.53
2018-10-08 | paphol |  | 24982.65 | 53.32 | 24973.81 | 50.27 | 24772 | 65
2018-11-05 | reet |  | 24956.71 | 42.71 | 24945.45 | 44.59 | 24850.24 | 41.41
2018-12-11 | eves/reet |  | 24916.4 | 38.27 | 24900 | 33.3 | 24798.97 | 33.8
2019-01-11 | eves |  |  |  | 24779.58 | 33.36 | 24818.82 | 37.98
2019-02-18 | paphol |  | 24796.3 | 33.78 | 24762.03 | 57.82 | 24815.18 | 34.51
2019-03-12 | eves
2019-03-13 | eves |  | 24823.39 | 53.85 | 24750.41 | 65.7 | 24834.66 | 82.67
2019-03-14 | paphol
2019-03-14 | paphol |  | 24832.9 | 43.25 | 24844.9 | 45.21 | 24811.4 | 50.46
2019-04-17 | eves |  | 24885.92 | 49.42 | 24886.82 | 48.56 | 24872.79 | 53.65
2019-05-23 | eves |  |  |  | 24889.08 | 46.39 | 24867.32 | 51.44
2019-06-21 | eves |  | 24937.25 | 50.075 | 24896.58 | 51.91 | 24875.31 | 46.58
2019-07-25 | eves |  | 24881.86 | 49.45 | 24880.31 | 51.91 | 24858.03 | 47.7
2019-08-26 | paphol |  | 24818.382 | 50.03 | 24825.708 | 48.55 | 24854.39 | 53.04
2019-10-02 | paphol |  | 24887.65 | 48.76 | 24894.84 | 50.59 | 24863.35 | 50.57
2019-12-03 | paphol |  | 24862.1 | 32.05 | 24859.93 | 32.259 | 24835.72 | 32.59
2020-01-06 | eves |  | 24878.2 | 47.46 | 24878.89 | 49.61
2020-02-04 | paphol |  | 24851.409 | 26.152 | 24850.9 | 34.57 | 24826 | 32.45
2020-02-18 | reet |  | 24877 | 46
2020-03-09 | reet |  | 24948 | 59 | 24927 | 46 | 24903 | 48
2020-04-23 | paphol |  | 24940.91 | 52 | 24931.54 | 49.61 | 24899.64 | 48.47
2020-05-25 | reet |  | 24928 | 35 | 24911 | 31 | 24888 | 29
2020-06-25 | reet |  | 24940 | 50 | 24939 | 51 | 24897 | 33
2020-07-29 | reet |  | 24912 | 44 | 24902 | 47 | 24877 | 37
2020-08-28 | paphol |  | 24911.75 | 51.69 | 24907.56 | 45.98 | 24879.22 | 49.9
2020-09-24 | reet |  | 24901 | 35 | 24899 | 34 | 24876 | 29
2020-10-07 | reet | new stylus | 24848 | 35 | 24859 | 50 | 24860 | 33
|  | average | 24883.8971 | 44.2396 | 24871.984 | 43.1928 | 24840.2032 | 42.0616
|  | std dev | 61.6523 | 9.7921 | 64.5498 | 9.3923 | 42.5688 | 11.9487