## Sheet1
Uncertainty budget for DektakXTA
in general
source of uncertainty | value +/- | shape of pdf | divisor | standard uncertainty | What is this uncertainty about?
uncertainty on calibration of standard step height | 0.15 µm for the 24.865 µm standard and 30 Å for the 9160 Å standard | Gaussian | 2 | depends on range. Systematic error | difference between what we expect that the standard step height is and what it truly is
uncertainty on QC measurements of the standard step height. We do calibration if outside limits | Treating the distribution of actual measured step heights as Gaussian distributions gives a "standard deviation" of 25 Å for the small range and 66 nm for the 65.5 micron range (after excluding the worst outliers). | Actual measurements of the standard step height are kinda Gaussian-distributed (actually a bit more square than a Gaussian and with some very distant outliers) and centered on 9161 Å for the 6.5 micron range and 24871 nm for the 65.5 micron range. Due to our QC procedure the distribution cuts off at 0.15 µm for the 24.865 µm standard (3 large ranges) and 137 Å for the 9160 Å standard (smallest range) | NA | can only be estimated.…could be a combination of systematic and random error | difference between the average value we measure for the standard step heigth with the Dektak and what we expect the standard step size to be
resolution | 1 Å, 10 Å, 80 Å or 160 Å (for ranges 65 kÅ, 655 kÅ, 5240 kÅ and 1 mm respectively) | uniform | sqrt 3 | depends on range. Random error | difference between the value the Dektak spits out and what it would spit out with better resolution…
standard uncertainty of mean (10 repeated readings) | 24 Å on 6.5 micron range range, 34 nm on larger ranges for the standard samples. Could be larger for other sample types | Gaussian | none | depends on range and on sample. Random error | difference between what we measure in one Dektak scan of some feature of interest and what we measure as the average of many scans
uncertainty on reading w 6.5 micron range
source of uncertainty | value +/- (nm) | shape of pdf | divisor | standard uncertainty (nm)
standard step height uncertainty | 30 | gauss | 2 | 15
QC uncertainty | 2.5 | distribution cut off at 0.15 µm for the 24.865 µm standard (3 large ranges) and 137 Å for the 9160 Å standard (smallest range) | 1 | 2.5
resolution | 0.1 | uniform | sqrt 3 | 0.17321
standard uncertainty of mean | 2.4 | gauss | none | 2.4
total uncertainty |  |  |  | 15.3961
95 % confidence int |  |  |  | 30.7922
uncertainty on reading w 65.5 micron range
source of uncertainty | value +/- (micron) | shape of pdf | divisor | standard uncertainty (microns)
standard step height uncertainty | 0.15 | gauss | 2 | 0.075
QC uncertainty | 0.07 | distribution cut off at 0.15 µm for the 24.865 µm standard (3 large ranges) and 137 Å for the 9160 Å standard (smallest range) | 1 | 0.07
resolution | 0.001 | uniform | sqrt 3 | 0.0017321
standard uncertainty of mean | 0.034 | gauss | none | 0.034
total uncertainty |  |  |  | 0.10809
95 % confidence int |  |  |  | 0.21619
uncertainty on reading w 524 micron range
source of uncertainty | value +/- (micron) | shape of pdf | divisor | standard uncertainty (microns)
standard step height uncertainty | 0.15 | gauss | 2 | 0.075
QC uncertainty | 0.07 | distribution cut off at 0.15 µm for the 24.865 µm standard (3 large ranges) and 137 Å for the 9160 Å standard (smallest range) | 1 | 0.07
resolution | 0.008 | uniform | sqrt 3 | 0.013856
standard uncertainty of mean | 0.03 | gauss | none | 0.03
total uncertainty |  |  |  | 0.10778
95 % confidence int |  |  |  | 0.21556
uncertainty on reading w 1 mm range
source of uncertainty | value +/- (micron) | shape of pdf | divisor | standard uncertainty (microns)
standard step height uncertainty | 0.15 | gauss | 2 | 0.075
QC uncertainty | 0.07 | distribution cut off at 0.15 µm for the 24.865 µm standard (3 large ranges) and 137 Å for the 9160 Å standard (smallest range) | 1 | 0.07
resolution | 0.18 | uniform | sqrt 3 | 0.31177
standard uncertainty of mean | 0.035 | gauss | none | 0.035
total uncertainty |  |  |  | 0.32257
95 % confidence int |  |  |  | 0.64514

## Sheet2
Uncertainty budget for Dektak 8
in general
source of uncertainty | value +/- | shape of pdf | divisor | standard uncertainty | What is this uncertainty about?
uncertainty on calibration of standard step height | 0.15 µm for the 24.865 µm standard and 300 Å for the 916 nm standard | gaussian | 2 | depends on range. Systematic error | difference between what we expect that the standard step height is and what it truly is
uncertainty on QC measurements of the standard step height. We do calibration if outside limits | actual measurements of the standard step height are kinda a gaussian centered on 9212 Å for the 6.5 micron range and 24860-24867 nm for the larger ranges with a "standard deviation" of 24 Å for the small range and 50-60 nm for the larger ranges. So for the smallest range the average value is actually off from the expected value of the standard by more than 2 "standard deviations" ! | distribution cut off at 0.15 µm for the 24.865 µm standard (3 large ranges) and 137 Å for the 9160 Å standard (smallest range) but this hardly matters as it's in the far far tails | NA | can only be estimated.…could be a combination of systematic and random error | difference between the average value we measure for the standard step heigth with the Dektak and what we expect the standard step size to be
resolution | 1 Å, 10 Å, 80 Å or 140 Å (for ranges 65 kÅ, 655 kÅ, 5240 kÅ and 1 mm respectively) | uniform | sqrt 3 | depends on range. Random error | difference between the value the Dektak spits out and what it would spit out with better resolution…
standard uncertainty of mean (10 repeated readings) | 20 Å on 6.5 micron range, 24-26 nm on larger ranges when measuring the standard step (can depend a lot on sample) | gaussian | none | depends on range and on sample. Random error | difference between what we measure in one Dektak scan of some feature of interest and what we measure as the average of many scans
uncertainty on reading w 6.5 micron range
source of uncertainty | value +/- (nm) | shape of pdf | divisor | standard uncertainty (nm)
standard step height uncertainty | 30 | gauss | 2 | 15
QC uncertainty | 3 random + 5 systematic | between gaussian and uniform | none | 8
resolution | 0.1 | uniform | sqrt 3 | 0.17321
standard uncertainty of mean | 2 | gauss | none | 2
total uncertainty |  |  |  | 17.1181
95 % confidence int |  |  |  | 34.2362
uncertainty on reading w 65.5 micron range
source of uncertainty | value +/- (micron) | shape of pdf | divisor | standard uncertainty (microns)
standard step height uncertainty | 0.15 | gauss | 2 | 0.075
QC uncertainty | 0.05 | between gaussian and uniform | none | 0.05
resolution | 0.001 | uniform | sqrt 3 | 0.0017321
standard uncertainty of mean | 0.024 | gauss | none | 0.024
total uncertainty |  |  |  | 0.093295
95 % confidence int |  |  |  | 0.18659
uncertainty on 200 micron reading w 262 micron range
source of uncertainty | value +/- (micron) | shape of pdf | divisor | standard uncertainty (microns)
standard step height uncertainty | 0.15 | gauss | 2 | 0.075
QC uncertainty | 0.06 | between gaussian and uniform | none | 0.06
resolution | 0.008 | uniform | sqrt 3 | 0.013856
standard uncertainty of mean | 0.025 | gauss | none | 0.025
total uncertainty |  |  |  | 0.10021
95 % confidence int |  |  |  | 0.20042
uncertainty onreading w 1 mm range
source of uncertainty | value +/- (micron) | shape of pdf | divisor | standard uncertainty (microns)
standard step height uncertainty | 0.15 | gauss | 2 | 0.075
QC uncertainty | 0.06 | between gaussian and uniform | none | 0.06
resolution | 0.16 | uniform | sqrt 3 | 0.27713
standard uncertainty of mean | 0.026 | gauss | none | 0.026
total uncertainty |  |  |  | 0.29445
95 % confidence int |  |  |  | 0.5889